## SCHEDULE OF LINE ITEMS
HTME VISN | STATION | EQUIPMENT PO | EQUIPMENT TYPE | QTY
VISN 1: VA New England Healthcare System | 402 AUGUSTA ME | 402B61001 | XR US | 1
VISN 1: VA New England Healthcare System | 523 JAMAICA PLAIN MA | 523B61010 | XR US | 2
VISN 10: VA Healthcare System of Ohio | 506 ANN ARBOR MI | 506B60039 | XR US | 2
VISN 10: VA Healthcare System of Ohio | 538 CHILLICOTHE OH | 538B60001 | XR US | 2
VISN 10: VA Healthcare System of Ohio | 552 DAYTON OH | 552B50018 | XR US | 4
VISN 10: VA Healthcare System of Ohio | 583 INDIANAPOLIS IN | 583B60039 | XR US | 3
VISN 10: VA Healthcare System of Ohio | 757 COLUMBUS OH | 757B60015 | XR US | 1
VISN 12: VA Great Lakes Health Care System | 578 HINES IL | 578B68013 | XR US | 1
VISN 12: VA Great Lakes Health Care System | 578 HINES IL | 578B68014 | XR US | 1
VISN 15: VA Heartland Network | 589 TOPEKA KS | 589B62002 | XR PORT C ARM | 1
VISN 15: VA Heartland Network | 589 COLUMBIA MO | 589B63016 | XR US | 1
VISN 15: VA Heartland Network | 589 COLUMBIA MO | 589B63022 | XR US | 2
VISN 16: South Central VA Health Care Network | 564 FAYETTEVILLE AR | 564B60002 | XR US | 5
VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | 580B60005 | XR US | 3
VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | 580B60006 | XR US | 2
VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | 580B60015 | XR US | 1
VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | 580B60019 | XR US | 7
VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | 580B60020 | XR US | 1
VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | 580B60022 | XR US | 2
VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | 580B60023 | XR US | 1
VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | 580B60025 | XR US | 1
VISN 16: South Central VA Health Care Network | 598 LITTLE ROCK AR | 598B60018 | XR MAMMO | 1
VISN 16: South Central VA Health Care Network | 598 LITTLE ROCK AR | 598B60020 | XR MAMMO | 1
VISN 17: VA Heart of Texas Health Care Network | 549 DALLAS TX | 549B60505 | XR US | 9
VISN 17: VA Heart of Texas Health Care Network | 671 SAN ANTONIO TX | 671B60042 | XR PORT C ARM | 1
VISN 17: VA Heart of Texas Health Care Network | 756 EL PASO TX | 756B65006 | XR US | 1
VISN 17: VA Heart of Texas Health Care Network | 756 EL PASO TX | 756B65007 | XR PORT C-ARM | 2
VISN 19: VA Rocky Mountain Network | 554 COLORADO SPRING CO | 554B60021 | XR-PORT C-ARM | 1
VISN 19: VA Rocky Mountain Network | 635 OKLAHOMA CITY OK | 635B60002 | XR US | 1
VISN 19: VA Rocky Mountain Network | 635 OKLAHOMA CITY OK | 635B60002 | XR US | 1
VISN 19: VA Rocky Mountain Network | 635 OKLAHOMA CITY OK | 635B60003 | XR PORT C ARM | 1
VISN 2: VA Healthcare Network Upstate New York | 526 BRONX NY | 526B60005 | XR US | 1
VISN 2: VA Healthcare Network Upstate New York | 632 NORTHPORT NY | 632B61006 | XR US | 1
VISN 20: VA Northwest Network | 648 PORTLAND OR | 648B63006 | XR US | 1
VISN 20: VA Northwest Network | 648 PORTLAND OR | 648B63007 | XR US | 1
VISN 20: VA Northwest Network | 668 SPOKANE WA | 668B63003 | XR US | 2
VISN 21: VA Sierra Pacific Network | 459 HONOLULU HI | 459B60029 | XR US | 3
VISN 21: VA Sierra Pacific Network | 570 FRESNO CA | 570B60010 | XR RAD | 1
VISN 21: VA Sierra Pacific Network | 570 FRESNO CA | 570B60011 | XR RAD | 1
VISN 21: VA Sierra Pacific Network | 570 FRESNO CA | 570B60012 | XR RAD | 1
VISN 21: VA Sierra Pacific Network | 612 MARINEZ CA | 612B69010 | XR US | 1
VISN 21: VA Sierra Pacific Network | 640 PALO ALTO CA | 640B64053 | XR US | 5
VISN 21: VA Sierra Pacific Network | 640 PALO ALTO CA | 640B64054 | XR US | 1
VISN 21: VA Sierra Pacific Network | 640 PALO ALTO CA | 640B64056 | XR US | 1
VISN 21: VA Sierra Pacific Network | 640 PALO ALTO CA | 640B64057 | XR PORT C ARM | 1
VISN 21: VA Sierra Pacific Network | 640 PALO ALTO CA | 640B64058 | XR PORT C ARM | 1
VISN 21: VA Sierra Pacific Network | 662 SAN FRANCISCO CA | 662B60014 | XR US | 3
VISN 21: VA Sierra Pacific Network | 662 SAN FRANCISCO CA | 662B60015 | XR US | 2
VISN 21: VA Sierra Pacific Network | 662 SAN FRANCISCO CA | 662B60025 | XR US | 1
VISN 21: VA Sierra Pacific Network | 662 SAN FRANCISCO CA | 662B60031 | XR PORT C ARM | 1
VISN 21: VA Sierra Pacific Network | 662 SAN FRANCISCO CA | 662B60032 | XR US | 4
VISN 22: VA Desert Pacific Healthcare Network | 644 PHOENIX AZ | 644B68023 | XR PORT C ARM | 2
VISN 22: VA Desert Pacific Healthcare Network | 644 PHOENIX AZ | 644B68024 | XR PORT C ARM | 1
VISN 22: VA Desert Pacific Healthcare Network | 644 PHOENIX AZ | 644B68025 | XR PORT C ARM | 1
VISN 22: VA Desert Pacific Healthcare Network | 664 SAN DIEGO CA | 664B60058 | XR US | 1
VISN 22: VA Desert Pacific Healthcare Network | 678 TUCSON AZ | 678B61031 | XR CT | 1
VISN 22: VA Desert Pacific Healthcare Network | 678 TUCSON AZ | 678B61032 | XR MRI | 1
VISN 22: VA Desert Pacific Healthcare Network | 691 LOS ANGELES CA | 691B60052 | XR PORT C ARM | 1
VISN 22: VA Desert Pacific Healthcare Network | 691 LOS ANGELES CA | 691B60053 | XR US | 2
VISN 23: VA Midwest Health Care Network | 499 SIOUX FALLS SD | 499B69000 | XR US | 1
VISN 23: VA Midwest Health Care Network | 499 IOWA CITY IA | 499B69004 | XR URO | 1
VISN 23: VA Midwest Health Care Network | 499 IOWA CITY IA | 499B69005 | XR US | 2
VISN 4: VA Stars & Stripes Healthcare Network | 460 WILMINGTON DE | 460B68008 | XR US | 1
VISN 4: VA Stars & Stripes Healthcare Network | 460 WILMINGTON DE | 460B68009 | XR US | 1
VISN 4: VA Stars & Stripes Healthcare Network | 595 LEBANON PA | 595B69013 | XR RAD | 1
VISN 4: VA Stars & Stripes Healthcare Network | 642 PHILADELPHIA PA | 642B69014 | XR US | 2
VISN 4: VA Stars & Stripes Healthcare Network | 642 PHILADELPHIA PA | 642B69016 | XR US | 1
VISN 4: VA Stars & Stripes Healthcare Network | 693 NEW ALLENTOWN PA | 693B63004 | XR MAMMO | 1
VISN 4: VA Stars & Stripes Healthcare Network | 693 WILKES BARRE PA | 693B63005 | XR RAD | 1
VISN 4: VA Stars & Stripes Healthcare Network | 693 WILKES BARRE PA | 693B63006 | XR US | 1
VISN 4: VA Stars & Stripes Healthcare Network | 693 WILKES BARRE PA | 693B63007 | XR RF | 1
VISN 4: VA Stars & Stripes Healthcare Network | 693 WILKES BARRE PA | 693B63008 | NM BONE DEN | 1
VISN 5: VA Capitol Health Care Network | 688 WASHINGTON DC | 688B61003 | XR US | 2
VISN 6: VA Mid-Atlantic Health Care Network | 558 DURHAM NC | 558B69009 | XR MAMMO | 1
VISN 6: VA Mid-Atlantic Health Care Network | 565 FAYETTEVILLE NC | 565B69015 | XR US | 1
VISN 6: VA Mid-Atlantic Health Care Network | 637 ASHEVILLE NC | 637B69003 | NM PET CT | 1
VISN 6: VA Mid-Atlantic Health Care Network | 659 SALISBURY NC | 659B69001 | XR US | 1
VISN 7: VA Southeast Health Care Network | 508 DECATUR GA | 508B62052 | XR US | 1
VISN 7: VA Southeast Health Care Network | 521 BIRMINGHAM AL | 521B62003 | XR US | 1
VISN 7: VA Southeast Health Care Network | 521 BIRMINGHAM AL | 521B62004 | XR PORT RAD | 1
VISN 7: VA Southeast Health Care Network | 521 BIRMINGHAM AL | 521B62007 | XR PORT RAD | 1
VISN 7: VA Southeast Health Care Network | 534 CHARLESTON SC | 534B62034 | NM BONE DEN | 1
VISN 7: VA Southeast Health Care Network | 544 COLUMBIA SC | 544B62014 | XR MAMMO | 1
VISN 7: VA Southeast Health Care Network | 557 DUBLIN GA | 557B62003 | XR PORT RAD | 2
VISN 7: VA Southeast Health Care Network | 557 DUBLIN GA | 557B62004 | XR US | 1
VISN 8: VA Sunshine Healthcare Network | 516 BAY PINES FL | 516B60008 | XR US | 2
VISN 8: VA Sunshine Healthcare Network | 546 MIAMI FL | 546B60023 | XR US | 1
VISN 8: VA Sunshine Healthcare Network | 546 MIAMI FL | 546B60024 | XR US | 2
VISN 8: VA Sunshine Healthcare Network | 546 MIAMI FL | 546B60025 | XR US | 2
VISN 8: VA Sunshine Healthcare Network | 573 TALLAHASSEE FL | 573B65025 | XR US | 1
VISN 8: VA Sunshine Healthcare Network | 573 TALLAHASSEE FL | 573B65026 | NM BONE DEN | 1
VISN 8: VA Sunshine Healthcare Network | 573 GAINESVILLE FL | 573B65030 | XR US | 2
VISN 8: VA Sunshine Healthcare Network | 573 GAINESVILLE FL | 573B65031 | XR US | 1
VISN 8: VA Sunshine Healthcare Network | 573 LAKE CITY FL | 573B65032 | XR PORT C ARM | 1
VISN 8: VA Sunshine Healthcare Network | 573 GAINESVILLE FL | 573B65034 | XR US | 1
VISN 8: VA Sunshine Healthcare Network | 573 LAKE CITY FL | 573B65035 | XR US | 2
VISN 8: VA Sunshine Healthcare Network | 573 GAINESVILLE FL | 573B65036 | XR US | 1
VISN 8: VA Sunshine Healthcare Network | 573 GAINESVILLE FL | 573B65037 | XR US | 1
VISN 8: VA Sunshine Healthcare Network | 573 TALLAHASSEE FL | 573B65038 | XR US | 2
VISN 8: VA Sunshine Healthcare Network | 672 SAN JUAN PR | 672FC6008 | XR US | 1
VISN 8: VA Sunshine Healthcare Network | 672 SAN JUAN PR | 672FC6012 | XR HYBRID | 1
VISN 8: VA Sunshine Healthcare Network | 675 ORLANDO FL | 675B67006 | XR MRI | 1
VISN 8: VA Sunshine Healthcare Network | 675 ORLANDO FL | 675B67007 | XR IR | 1
VISN 8: VA Sunshine Healthcare Network | 675 ORLANDO FL | 675B67012 | XR US | 2
VISN 8: VA Sunshine Healthcare Network | 675 ORLANDO FL | 675B67014 | XR US | 2
VISN 8: VA Sunshine Healthcare Network | 675 ORLANDO FL | 675B67015 | XR US | 1
VISN 8: VA Sunshine Healthcare Network | 675 ORLANDO FL | 675B67016 | XR PORT C ARM | 1
VISN 8: VA Sunshine Healthcare Network | 675 ORLANDO FL | 675B67017 | XR MRI | 1
VISN 9: VA Mid South Healthcare Network | 621 MOUNTAIN HOME TN | 621B60002 | XR US | 2
VISN 9: VA Mid South Healthcare Network | 626 CHATTANOOGA TN | 626B64028 | XR US | 1

## QA TEMPLATE
EQUIPMENT PO | VISN | STATION | EQUIPMENT TYPE | VENDOR EQUIPMENT QUESTIONS | GOVERNMENT EQUIPMENT RESPONSE
460B68009 | VISN 4: VA Stars & Stripes Healthcare Network | 460 WILMINGTON DE | XR US | Request item 1.k. Maximum equipment weightbe increased from 87 lbs to 141 lbs. Under Section 6. Analysis Section - do you require any automated software for IUD Placement? | This Feature is not a requirement
460B68009 | VISN 4: VA Stars & Stripes Healthcare Network | 460 WILMINGTON DE | XR US | 1.k. Can the maximum equipment weight be raised to 126 lbs to increasee competition? | YES
460B68009 | VISN 4: VA Stars & Stripes Healthcare Network | 460 WILMINGTON DE | XR US | 7.a. Can the PureWave feature be removed from the requirement as it appears to be brand-specific? | YES
499B69000 | VISN 23: VA Midwest Health Care Network | 499 SIOUX FALLS SD | XR US | VISN 23: VA Sious Falls What surgical procedures are to be performed now or in the future or is this soley for Transrectal and Transperineal Prostate procedures? If required for surgery, additional software applications and probes are required. Image Fusion Software. Please confirm what procedures will be performed utilizing Fusion. If Fusion is required, will it be needed for TP, TR, or both? If Fusion is required, will it be required to be freehand or stabilized? | System solely used for TR and TP procedures. No Image fusion software.
499B69004 | VISN 23: VA Midwest Health Care Network | 499 IOWA CITY IA | XR URO | Line 31K - Would changing the minimum Xray monitor size to 19 inches be considered a clinically equivalent or substantially complaint alternative to the specified 27 inch monitor size? | Yes
499B69004 | VISN 23: VA Midwest Health Care Network | 499 IOWA CITY IA | XR URO | Line 32L-Would changing the minimum video monitor size to 19 inches be considered a clinically equivalent or substantially compliant alternative to the specified 27 inch monitor size? | Yes
499B69004 | VISN 23: VA Midwest Health Care Network | 499 IOWA CITY IA | XR URO | line 28 h. Please verify the vertical travel spec requested. Can the vertical travel spec be changed to 20.7" for fair competition. | Yes
499B69005 | VISN 23: VA Midwest Health Care Network | 499 IOWA CITY IA | XR US | Line 22C- Please advise if the control panel height range from the floor in 57.5-67.5 inches is correct requirement? | Yes
499B69005 | VISN 23: VA Midwest Health Care Network | 499 IOWA CITY IA | XR US | Line 24E- Is the minimum storage of 2TB a combination of the operating software and patient information storage? | Yes
499B69005 | VISN 23: VA Midwest Health Care Network | 499 IOWA CITY IA | XR US | Line 58I - Please explain Image editing feature and how it applies to the sonographer workflow? | It allows us to optimize image after grabbed, improves sonographer productivity and exam quality
499B69005 | VISN 23: VA Midwest Health Care Network | 499 IOWA CITY IA | XR US | Request item 1.g. Maximum Weight be increased to 265 lbs. | Yes
506B60039 | VISN 10: VA Healthcare System of Ohio | 506 ANN ARBOR MI | XR US | Request 1.a. Minimum screen size reduced to 14 in. Request 2.m. Other scanning modes PIV moved to Added Value. | 1.a: Minimum screen size remains at 21 inches, as reducing it to 14 inches would significantly compromise the clinical users' needs. Clinical staff require a larger display to support effective treatment and diagnosis, which ultimately benefits patient care. • 2.m: The requirement for "Other scanning modes" is retained as a technical requirement, while reference to a specific scanning mode has been removed to allow vendors flexibility in meeting the clinical need. The underlying clinical justification is that access to multiple scanning modes is essential for staff to optimize image acquisition and adapt to diverse diagnostic and therapeutic scenarios, thereby supporting high standards of patient care and enabling effective diagnosis and treatment.
508B62052 | VISN 7: VA Southeast Health Care Network | 508 DECATUR GA | XR US | Missing technical specifications and matrix forms | Technical specifications and compliance matrix added
508B62052 | VISN 7: VA Southeast Health Care Network | 508 DECATUR GA | XR US | Benchmark missing | Technical specifications and compliance matrix added
516B60008 | VISN 8: VA Sunshine Healthcare Network | 516 BAY PINES FL | XR US | Request item 1.a. Minimum screen size be revised to 19 in. Request item 9.g. Integration to Dose Tracking be removed. | Tech Specs updated
521B62003 | VISN 7: VA Southeast Health Care Network | 521 BIRMINGHAM AL | XR US | A battery life of 60 minutes is required. The system will be used for liver resection and vascular surgical cases. Is a battery necessary since the system will not leave the surgical suite? | Yes, the battery is necessary for transport from the room to the storage area where it will remain plugged in.
521B62003 | VISN 7: VA Southeast Health Care Network | 521 BIRMINGHAM AL | XR US | Can the minimum monitor size be updated to 19 inches to increase competition? Can the minimum active screen be updated to 19 inches to increase competition? Can the minimum transducer ports be updated to 4 to increase competition? Can you please confirm functionality of the system as the description details liver resection and vascular procedures but lists obstetric, Mitral valve navigator, and Cardiac 3D Quantification. Can the b. 4-way lap C9150LA frequency range be clarified. Image Fusion Software. Please confirm what procedures will be performed utilizing Fusion. If Fusion is required, will it be needed for TP, TR, or both? If Fusion is required, will it be required to be freehand or stabilized? | No, the minimum monitor size of 23" is requested by the clinical team for ease of viewing No, the minimum active screen of 23" is requested by the clinical team for ease of viewing No, the minimum transducer cannot be updated to 4 Ultrasound will be used for urology and general surgery Can the b. 4-way lap C9150LA frequency range be clarified? Frequency range: 5.0 MHz (low) – 13.0 MHz (high). See updated Technical specifications and compliance matrix Fusion software will not be used
521B62004 | VISN 7: VA Southeast Health Care Network | 521 BIRMINGHAM AL | XR PORT RAD | Detector Pixel Size • Would a detector with a 148 μm pixel size be considered a clinically equivalent or substantially compliant alternative to the specified 140 μm maximum pixel size requirement? Display Monitor Size • Would an integrated 17-inch display monitor be considered substantially compliant with the 21.5-inch minimum monitor size requirement if it provides equivalent workflow functionality and image review capabilities? Detector Weight Capacity • Would a detector with a 661 lb recumbent patient weight capacity be considered substantially compliant with the specified 881 lb distributed load capacity requirement, and if so, what clinical justification would be needed to support equivalency? Generator kVp Range • Would a system with a maximum generator output of 133 kVp be considered substantially compliant with the specified 150 kVp maximum generator requirement, provided all clinical imaging requirements can be met? Display at Tube Head • Is a dedicated monitor at the tube head a mandatory requirement, or would systems with a centralized touchscreen workstation providing equivalent positioning and workflow functionality be considered acceptable? | A detector with 148 um pixel size would be considered clinically equivalent to 140 um pixel size. Integrated 17-inch display monitor would be considered compliant with the 21.5 if the workflow functionality and image review capabilities are comparable. A detector with 661lbs recumbent patient weight capacity would be compliant to 881lbs distributed load capacity. The clinical need is for the bariatric patients. Maximum generator of 133kVp would be considered compliant with the 150 kVp maximum generator requirement. Display monitor on tube head would be a preferred option but not required. A centralized touchscreen workstation providing equivalent positioning and workflow would be considered an alternative option but not required. See updated Technical specifications and compliance matrix
521B62004 | VISN 7: VA Southeast Health Care Network | 521 BIRMINGHAM AL | XR PORT RAD | 1.a. Can you please change the generator power to 30kW for more open competition? 3.c. Please move the free release handles to value adds. 3.d. Please move the positioning light to value adds. 3.g. Please move ability to move portable from tube head to value adds. 3.p. Please move display monitor on tube head to value adds. 3.r. Please move SID positioning aides to value adds. 3.s. Please move tube and detector angle alignment software to value adds. | Generator power was changed to 30kW on the updated technical specifications and compliance matrix Free release handles were moved to Added Value on the Tech Specs and compliance matrix. Positioning light is required. The ability to move portable from tube head is required. Display monitor on tube head would be a preferred option but not required. SID positioning aides are required. Tube and detector angle alignment software would be a preferred option but not required. This is moved to Added Value on the Tech Specs and compliance matrix.
521B62007 | VISN 7: VA Southeast Health Care Network | 521 BIRMINGHAM AL | XR PORT RAD | 2.c. Please confirm need or remove need for navigation software. Usually this is used in IR or for procedures. 2.d. Please confirm need or remove need for image fusion software. | Navigation software is needed because the IR physician use ultrasound for procedures such as injections and drainages. Image fusion software can be an option but not required. This was moved to Added Value on the Tech Specs and compliance matrix.
526B60005 | VISN 2: VA Healthcare Network Upstate New York | 526 BRONX NY | XR US | Image Fusion Software. Please confirm what procedures will be performed utilizing Fusion. If Fusion is required, will it be needed for TP, TR, or both? If Fusion is required, will it be required to be freehand or stabilized? Can 1. Clinical Training d. Technologists who complete the clinical applications training shall receive continuing education credits (CMEs) be added value? | 1. Fusion is required for both TP and TR. 2. Fusion should be freehand and stabilized. 3. Fine for it to be added value.
538B60001 | VISN 10: VA Healthcare System of Ohio | 538 CHILLICOTHE OH | XR US | 2.b. Please confirm whether 3D freehand is needed. Request to have this moved to the value-add section or removed, as it is not needed. | Yes, Please see revised Tech Specs and Compliance Matrix for changes.
538B60001 | VISN 10: VA Healthcare System of Ohio | 538 CHILLICOTHE OH | XR US | 2.c. Please confirm if navigation software is needed. If not, please remove or move to value-add. | Yes, Please see revised Tech Specs and Compliance Matrix for changes.
538B60001 | VISN 10: VA Healthcare System of Ohio | 538 CHILLICOTHE OH | XR US | 2.d. Please confirm if image fusion software is needed. If not, please remove or move to value-add. | Yes, Please see revised Tech Specs and Compliance Matrix for changes.
544B62014 | VISN 7: VA Southeast Health Care Network | 544 COLUMBIA SC | XR MAMMO | Will site also be trading in an Affirm Upright biopsy system? If so, please provide the SN? It begins with 820 | We encourage and expect vendors to deinstall + trade-in old systems. Affirm Biopsy SN is 82011151298
544B62014 | VISN 7: VA Southeast Health Care Network | 544 COLUMBIA SC | XR MAMMO | 1e. SID 70cm - to allow for fair competition please update to 65-70. h. Max system weight 900lbs - to allow for fair competition please update to 925 to 1. What vacuum assisted device will you use for biopsy? 2. Can the window of the weight be widened to maximum weight to 925 to be more vendor neutral? | Ok to set SID requirement to 65-70cm, updated on latest technical specifications and compliance matrix OK to set weight requirement to 925lb, updated on latest technical specifications and compliance matrix No, we are requiring a self-contained vacuum device/system. To provide otherwise will require site prep, which is not desired.
546B60023 | VISN 8: VA Sunshine Healthcare Network | 546 MIAMI FL | XR US | Line 27 - Requested height range only says 30". Is this from the floor? Our spec sheet does not specify height from floor. Line 31 - Are they requesting a portable unit? These specifications are quite small. Line 66 - CV systems don't offer Small Parts. Will this be necessary? | Line 27- 1h. Removed specification. Original ask was total adjustable range for the control panel. Line 31- 1l. Not requesting portable unit. Corrected equipment dimensions to 74in x 25 in x 36in Line 66- 6d. Removed specification.
546B60024 | VISN 8: VA Sunshine Healthcare Network | 546 MIAMI FL | XR US | Request item 1.a. and 1.b be reduced to 14 in. Request item 1.i. Minimum battery life be reduced to 60 min. | Tech Specs updated
546B60025 | VISN 8: VA Sunshine Healthcare Network | 546 MIAMI FL | XR US | Image Fusion Software. Please confirm what procedures will be performed utilizing Fusion. If Fusion is required, will it be needed for TP, TR, or both? If Fusion is required, will it be required to be freehand or stabilized? | This requirement is not on the tech specs. No changes made.
549B60505 | VISN 17: VA Heart of Texas Health Care Network | 549 DALLAS TX | XR US | Line 35B Please explain what procedures you will be using for Three-dimensional (3D) freehand scanning mode | Evaluation of valvular stenosis and regurgitation, evaluation of vegetations, evaluation of Watchman and other structual devices
549B60505 | VISN 17: VA Heart of Texas Health Care Network | 549 DALLAS TX | XR US | Line 68J Please explain what procedures you will be using for shear wave and stain-based elastography | LV systolic function in patients on chemotherapy, infiltrative heart disease evalation
549B60505 | VISN 17: VA Heart of Texas Health Care Network | 549 DALLAS TX | XR US | Lines 87A-91E Please clarify how you would like the quantities of the transducers reflected on the 9 units | Each unit will have a 2D/3D transthoracic probe and pedhoff probe. 2D/3D transesophageal probes will be used on 4-5 machines.
549B60505 | VISN 17: VA Heart of Texas Health Care Network | 549 DALLAS TX | XR US | For items 5q and 7.3. The solicitation calls for 9 CVUS systems, but only one ICE probe. Would you like for ICE software to be loaded on all 9 CVUS systems or just on 1 system?? | One system will need Versight capabilities, with a PIM cable to unlock these capabilities.
549B60505 | VISN 17: VA Heart of Texas Health Care Network | 549 DALLAS TX | XR US | 5.j – Shear Wave and Strain-Based Elastography: Can you please clarify the intended cardiovascular clinical application for shear wave and strain-based elastography within this requirement? Is this referring to cardiac strain imaging, or is tissue elastography specifically required? 6.a – Abdominal Analysis Package: The intended clinical applications outlined in the solicitation are TTE, TEE, stress echocardiography, Doppler echocardiography, and other cardiovascular cases. Can you please clarify the intended cardiovascular clinical application for the abdominal analysis package and whether this functionality is required for the systems being procured? 7.a, 7.d & 7.e – Transducer Quantities: The solicitation requests nine ultrasound systems; however, the transthoracic echo and Doppler transducers are each listed at a quantity of 12. As well as the Intracardiac is only listed as 1. Can you please confirm the required quantities for these transducers to ensure the proposal includes the quantities intended by the VA? | 5j. Yes , this is is supposed to be for LV/RV/LA strain. 6a. The abdominal package has been removed as a requirement. 7a.d.e. We need to pair each 2D/3D tranthoracic probe with each of the 9 machines. We will use the extra 3 2D/3D probes as backups if one of the 9 fail. We will pair each of the Doppler (pedoff) probes with each of the 9 machines. These probes aren't used as frequently, so we only need 9 of these rather than 12. 7.b.c [NOTE] For the TEE probes, it is possible to lower the requirement to 6 of the regular thickness probes (8-2 MHz) and 3 of the thinner probes (11-4 MHz) (rather than 8 and 4). The requirement has been adjusted.
552B50018 | VISN 10: VA Healthcare System of Ohio | 552 DAYTON OH | XR US | 2.b. Please remove navigation software; this is not applicable for this care area. | Yes, Please see revised Tech Specs and Compliance Matrix for changes.
552B50018 | VISN 10: VA Healthcare System of Ohio | 552 DAYTON OH | XR US | 2.c. Please remove image fusion software; this is not applicable for this care area. | Yes, Please see revised Tech Specs and Compliance Matrix for changes.
554B60021 | VISN 19: VA Rocky Mountain Network | 554 COLORADO SPRING CO | XR-PORT C-ARM | line 26 f. Can orbital rotation spec be changed to 120 degrees for fair competition? | Changing to 120 is acceptable
554B60021 | VISN 19: VA Rocky Mountain Network | 554 COLORADO SPRING CO | XR-PORT C-ARM | line 27 g. Can pivot rotation spec be changed to 380 degrees for fair competition? | Changing to 380 is acceptable
554B60021 | VISN 19: VA Rocky Mountain Network | 554 COLORADO SPRING CO | XR-PORT C-ARM | line 31 k. Can minimum C-arm depth spec be changed to 18"for fair competition? | Changing to 18" is acceptable
554B60021 | VISN 19: VA Rocky Mountain Network | 554 COLORADO SPRING CO | XR-PORT C-ARM | line 34 n. Can image matrix size spec be changed to 1536x1532 pixels for fair competition? | Changing the spec is acceptable
554B60021 | VISN 19: VA Rocky Mountain Network | 554 COLORADO SPRING CO | XR-PORT C-ARM | line 47 j. Can the "rainbow" spec be changed to 120 degrees for fair competion? | Current spec is 180 deg so this would be a 33% reduction and could make it more difficult to use. Do not concur.
554B60021 | VISN 19: VA Rocky Mountain Network | 554 COLORADO SPRING CO | XR-PORT C-ARM | line 51 n. This is typically a full size c-arm feature, not available on mini, can this be moved to value add for fair competition? | Removed
557B62003 | VISN 7: VA Southeast Health Care Network | 557 DUBLIN GA | XR PORT RAD | Detector Pixel Size • Would a detector with a 148 μm pixel size be considered a clinically equivalent or substantially compliant alternative to the specified 140 μm maximum pixel size requirement? Display Monitor Size • Would an integrated 17-inch display monitor be considered substantially compliant with the 21.5-inch minimum monitor size requirement if it provides equivalent workflow functionality and image review capabilities? Detector Weight Capacity • Would a detector with a 661 lb recumbent patient weight capacity be considered substantially compliant with the specified 881 lb distributed load capacity requirement, and if so, what clinical justification would be needed to support equivalency? Generator kVp Range • Would a system with a maximum generator output of 133 kVp be considered substantially compliant with the specified 150 kVp maximum generator requirement, provided all clinical imaging requirements can be met? Display at Tube Head • Is a dedicated monitor at the tube head a mandatory requirement, or would systems with a centralized touchscreen workstation providing equivalent positioning and workflow functionality be considered acceptable? | A detector with 148 um pixel size would be considered clinically equivalent to 140 um pixel size. Integrated 17-inch display monitor would be considered compliant with the 21.5 if the workflow functionality and image review capabilities are comparable. A detector with 661lbs recumbent patient weight capacity would be compliant to 881lbs distributed load capacity. Given our facility’s patient demographics, the highest patient weight encountered in the past three years has been less than 661 lbs. The portable detector would include safeguards and workflows to prevent use with patients above the manufacturer-rated capacity. The clinical risk is minimal and does not compromise patient safety or care. The 661 lb detector meets our operational needs, and alternative imaging arrangements are in place for higher-weight patients. Maximum generator of 133kVp would be considered compliant with the 150 kVp maximum generator requirement. There is no anticipated impact on clinical care or diagnostic image quality. Exams requiring higher kVp are performed on fixed radiography systems. Thus, the proposed system with 133 kVp maximum output meets our operational and clinical requirements. A dedicated display at the tube head is not a mandatory requirement. Systems that provide a centralized touchscreen workstation—offering equivalent functionality in terms of positioning guidance, workflow integration, and easy access to controls. A tube head display is not required, but having convenient access to critical workflow tools near the imaging area is very helpful for technologist ease of use, and the selection should consider how each detector system supports these needs. See updated technical specifications and compliance matrix
557B62003 | VISN 7: VA Southeast Health Care Network | 557 DUBLIN GA | XR PORT RAD | 1.a. Can you please change the generator power to 30kW for more open competition? 3.c. Please move the free release handles to value adds. 3.d. Please move the positioning light to value adds. 3.g. Please move ability to move portable from tube head to value adds. 3.p. Please move display monitor on tube head to value adds. 3.r. Please move SID positioning aides to value adds. 3.s. Please move tube and detector angle alignment software to value adds. | Yes, please see updated technical specifications and compliance matrix.
557B62004 | VISN 7: VA Southeast Health Care Network | 557 DUBLIN GA | XR US | Missing transducers on technical specification form | Yes, please see updated technical specifications and compliance matrix.
557B62004 | VISN 7: VA Southeast Health Care Network | 557 DUBLIN GA | XR US | Line 61, 70, and 73: These are not cardiovascular clinical options, should they be removed? Line 81: Can this be changed to a Linear Array Probe (as 9L Array Probe is a single Vendor specific probe)? Line 82: Can this be changed to a Convex Array Probe (as C1-6 Convex Array Probe is a single vendor specific probe)? Line 83: Can this be changed to a Phased Array Probe (as 4VC-D Probe is a single Vendor specific probe)? | Yes, please see updated technical specifications and compliance matrix.
558B69009 | VISN 6: VA Mid-Atlantic Health Care Network | 558 DURHAM NC | XR MAMMO | Please provide the correct system trade-in SN. It will start with 810 or SDM. | Updated
558B69009 | VISN 6: VA Mid-Atlantic Health Care Network | 558 DURHAM NC | XR MAMMO | 1.q. Can the software license be changed from "C-View Software License" to synthetic software license to allow for vendor neutrality? 2.c. Curved/Flexed Paddle Options: Curved paddle is vendor specific. Please remove for open competition. 3.d. Comfort-focused paddles: This is vendor specific question. Request to be restated as "Patient comfort aides available at gantry. 4. Trade In Serial Number. Please validate serial number on Trade-in. It appears the serial number is incorrect "6525-433278" and should start with "8" versus "6". | 1q. Yes. Call it synthetic software license 2c. Removed from technical specification form 3d. Removed from Technical specification form and updated to Patient comfort aides available at gantry 4. SDM131700391
565B69015 | VISN 6: VA Mid-Atlantic Health Care Network | 565 FAYETTEVILLE NC | XR US | Request item 1.e. be reduced to 60 min. | Reduced to 20 min
565B69015 | VISN 6: VA Mid-Atlantic Health Care Network | 565 FAYETTEVILLE NC | XR US | 1.e. Please remove minimum battery life or change to 20min. No vendor can support 240min on a radiology unit. 5.n. Please remove ECG requirement as there is no stated need for a cardiac (phased array) probe. | 1.e reduced to 20 min 5.n. Requirement removed
573B65025 | VISN 8: VA Sunshine Healthcare Network | 573 TALLAHASSEE FL | XR US | The technical specificaton & matrix forms is listed as Tallahasse, FL; however when you open them it is listed as Gainville, Fl. Can you confirm site of service address. | Gainesville, FL.
573B65025 | VISN 8: VA Sunshine Healthcare Network | 573 TALLAHASSEE FL | XR US | 1.i. Please remove minimum battery life or change to 20min. | Tech Specs updated
573B65030 | VISN 8: VA Sunshine Healthcare Network | 573 GAINESVILLE FL | XR US | Request item 1.f. Maximum Equipment weight be increased to 155 lbs. | Tech Specs updated
573B65030 | VISN 8: VA Sunshine Healthcare Network | 573 GAINESVILLE FL | XR US | 2.c. Please confirm need or remove need for navigation software. Usually this is used in IR or for procedures- not seen in ICU or critical care. 2.d. Please confirm need or remove need for image fusion software. | Tech Specs updated
573B65031 | VISN 8: VA Sunshine Healthcare Network | 573 GAINESVILLE FL | XR US | Do you need a D2 non imaging probe? | Yes, D2 is required
573B65031 | VISN 8: VA Sunshine Healthcare Network | 573 GAINESVILLE FL | XR US | For 2.c. Image Fusion - Can you verify if Imag Fusion Software - combining real-time Ultrasound Images with CT, MRI, or PET Scans - is required? | Tech Specs updated
573B65032 | VISN 8: VA Sunshine Healthcare Network | 573 LAKE CITY FL | XR PORT C ARM | Please clarify minimum horizontal travel, no c-arm meets 104cm | Tech Specs updated
573B65032 | VISN 8: VA Sunshine Healthcare Network | 573 LAKE CITY FL | XR PORT C ARM | Please adjust minimum c-arm depth to 27" to allow for fair and open competition | Tech Specs updated
573B65032 | VISN 8: VA Sunshine Healthcare Network | 573 LAKE CITY FL | XR PORT C ARM | Please adjust maximum weight for fair and open competition to 920 lbs | Tech Specs updated
573B65034 | VISN 8: VA Sunshine Healthcare Network | 573 GAINESVILLE FL | XR US | Request item 1.f. Control Panel Height be revised to 36 in. Request item 1.i. Minimum battery life be revised to 30 min. Request item 1.k. Maximum eqiupment weight be increased to 250. Request item 8.a, 8.b, and 8.c. be removed as not typically done in Cardiology. | Tech Specs updated
573B65034 | VISN 8: VA Sunshine Healthcare Network | 573 GAINESVILLE FL | XR US | Line 25: The height is listed as 575 inches. Based on the previous matrix, should this be corrected to 10 inches? Lines 74, 75, and 76: These do not appear to be cardiovascular clinical options. Can these lines be removed from the matrix? Line 91: This does not appear to be a cardiovascular option. Can this line be removed? Line 94: This does not appear to be a cardiovascular option. Can this line be removed? | Tech Specs updated
573B65035 | VISN 8: VA Sunshine Healthcare Network | 573 LAKE CITY FL | XR US | Request item 1.g. Minimum image Storage of 6TB be reduced to 1TB. | Tech Specs updated
573B65035 | VISN 8: VA Sunshine Healthcare Network | 573 LAKE CITY FL | XR US | 1.h. Please remove minimum battery life or change to 20min to allow for full and fair competition. 1.g. Please change minimum image storage to 1TB to allow for full and fair competition. | Tech Specs updated
573B65036 | VISN 8: VA Sunshine Healthcare Network | 573 GAINESVILLE FL | XR US | Request item 1.h. Minimum battery life be revised to 30 min. Request item 1.j. Maximum eqiupment weight be increased to 250. | Tech Specs updated
573B65036 | VISN 8: VA Sunshine Healthcare Network | 573 GAINESVILLE FL | XR US | Line 25: The height is listed as 575 inches. Based on the previous matrix, should this be corrected to 10 inches? Lines 74, 75, and 76: These do not appear to be cardiovascular clinical options. Can these lines be removed from the matrix? Line 91: This does not appear to be a cardiovascular option. Can this line be removed? Line 94: This does not appear to be a cardiovascular option. Can this line be removed? | Line 25: Refers to minimum ports - 3 Lines 74,75, and 76: Removed Line 91: The original matrix indicates line 91 is OEM supported operating systems and must remain. Line 94: The original matrix indicates line 94 is PACS compatbility and must remain.
573B65037 | VISN 8: VA Sunshine Healthcare Network | 573 GAINESVILLE FL | XR US | Request item 1.a. and 1.b. be revised to 24 in. Request item 8.b, and 8.c. be removed as not typically done in Gastroenterology. | Tech Specs updated
573B65038 | VISN 8: VA Sunshine Healthcare Network | 573 TALLAHASSEE FL | XR US | 1.i. Please remove minimum battery life or change to 20min. 5.m. Please remove phantom requirement or move to value-add section. This is usually a third party purchase. | Tech Specs updated
578B68013 | VISN 12: VA Great Lakes Health Care System | 578 HINES IL | XR US | Can you verify that 2.c Navigation and 2.d. Image Fusion will be used in this department? Transducers - can you verify if all of these transducers on the matrix are required? | 2c&d-No Transducer-We will not need all of the transducers. We will just need the ML4-20-D probe and L6-24-D Array probe for mammography. We could also any standard probes that come with the machine.
578B68013 | VISN 12: VA Great Lakes Health Care System | 578 HINES IL | XR US | G. Is there a trade in? | no
578B68013 | VISN 12: VA Great Lakes Health Care System | 578 HINES IL | XR US | Considering the overview notes that the machine will be used exclusively for Breast ultrasound imaging, can you please clarify/confirm the following: 1. Do you still require Navigation Software and Image Fusion Software (under Section 2 - Scanning Modes) 2. Do you require Analysis packages besides breast if the system is used exclusively for breast ultrasound (i.e. Abdominal, Pelvic, Urology, Obstetric, Cardiovascular, etc.) 3. The transducer section lists 10 various transducers. a. Considering the system will be used exclusively for breast ultrasound, do you want only transducers specific for breast imaging? i. If yes, which transducers? ii. If no, questions: 1. There is a quantity of two (2) 4-15 MHz transducers noted- Do you want a quantity of two? 2. There is a quantity of two (2) 2-9 MHz transducers noted - Do you want a quantity of two? | 1. No 2. No 3i. Transducers for breast imaging ML4-20-D Probe L6-24-D Linear Array Probe Any other probes that may come with the machine. 3ii - In the Tech Specs, it looks like these are referring to ports and not transducers? Please clarify.
578B68014 | VISN 12: VA Great Lakes Health Care System | 578 HINES IL | XR US | Request 4.a. a keyboard on the display be required in lieu of the Physical keyboard? | Yes
578B68014 | VISN 12: VA Great Lakes Health Care System | 578 HINES IL | XR US | 2.b. Please remove navigation software; this is not applicable for this care area. | updated tech specs
578B68014 | VISN 12: VA Great Lakes Health Care System | 578 HINES IL | XR US | 2.c. Please remove image fusion software; this is not applicable for this care area. | updated tech specs
578B68014 | VISN 12: VA Great Lakes Health Care System | 578 HINES IL | XR US | 5.m. Please remove the phantom requirement or move it to the value-add section. This is usually a third-party purchase. | updated tech specs
580B60005 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | Request item 1.m. Maximum equipment weight be increased to 265 lbs. | Maximum weight changed to 265 lbs.
580B60005 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | Please confirm the care area for this tech spec. Only two probes are selected (high-frequency small parts and high-frequency linear), and trade-ins are point-of-care systems. Is this for radiology or point-of-care (vascular access)? | Radiology - Intervential Radiology - PICC Team
580B60005 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | 2.b. Please confirm whether 3D freehand is needed for use at the CBOC. Request to have this moved to the value-add section or removed, as it is not needed. | This purchase is needed for the Main VA Hospital in Houston
580B60005 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | 2.c. Please confirm if navigation software is needed. If not, please remove or move to value-add. | Not needed
580B60005 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | 2.d. Please confirm if image fusion software is needed. If not, please remove or move to value-add. | Not needed
580B60005 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | 5.m. Please remove the phantom requirement or move it to the value-add section. This is usually a third-party purchase. | Added to Added Value section
580B60005 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | 6. Only two selections have been made: (a) Procedure and (e) Vascular. Please confirm whether more packages need to be included. | Confirmed. Only procedural and vasuclar are needed.
580B60005 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | 7. Please confirm these are the only two probes required for radiology. | Yes, only those 2 probes are needed.
580B60006 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | The technical specs read 580-B70027 but the Compliance Matrix reads 580-B0006, along with the spreadsheet. Please clarify which number is correct. | Correct number is 580B60006 - Technical specifications and compliance matrix added
580B60006 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | Column A does not include this Equipment PO - 580B6006 in order to pre populate the VISN, Station and Equipment type. Missing specs benchmark | Correct number is 580B60006 - Technical specifications and compliance matrix added
580B60006 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | Benchmark Missing | Correct number is 580B60006 - Technical specifications and compliance matrix added
580B60006 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | No tech spec or matrix found. | Correct number is 580B60006 - Technical specifications and compliance matrix added
580B60006 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | Tech spec provided but not on schedule of line items. This appears to be a 2027 item. | Correct number is 580B60006 - Technical specifications and compliance matrix added
580B60006 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | 3.b. Please confirm whether 3D freehand is needed. Request to have this moved to the value-add section or removed, as it is not needed. | Yes, it is needed
580B60006 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | 3.c. Please confirm if navigation software is needed. If not, please remove or move to value-add. | Not needed
580B60006 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | 3.d. Please confirm if image fusion software is needed. If not, please remove or move to value-add. | Not needed
580B60006 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | 6.j. Please confirm if shear-wave and/or strain-based elastography clinically needed. I don't believe any of the Houston sites are doing shear-wave. If not, please remove or move to the value-add section. | Not needed
580B60006 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | 6.p. Please remove the phantom requirement or move it to the value-add section. This is usually a third-party purchase. | Moved to added value section
580B60006 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | 8. Please provide more clarity on the transducers (linear, curved, microconvex); many of these have what appear to be brand features included. | Curvilinear probe 5-1 MHZ Linear probes with the following frequency range 12-3 MHz,18-4 MHz & 26-8 MHz Endocavitary probe 10-3 MHz
580B60006 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | 9.a-g. Please confirm whether the fusion-targeted biopsy system, along with associated lines, is needed; if not, remove it or move it to the value-add section. | Not needed
580B60006 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | g. Please confirm trade-in; three systems are listed to be traded in, but only two are requested. | 2 trade ins confirmed
580B60015 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | Request items 8.b., 8.c., 8.e, and 8.f. be removed as this is not typically used in this department. | needed by clinicians
580B60015 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | 2.c. Please confirm if navigation software is needed. If not, please remove or move to value-add. | needed by clinicians
580B60015 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | 2.d. Please confirm if image fusion software is needed. If not, please remove or move to value-add. | It is needed.
580B60015 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | 5.l. Please remove the phantom requirement or move it to the value-add section. This is usually a third-party purchase. | Moved to added value section
580B60015 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | 7. Please provide transducer clarification; the section appears to be focused on one vendor without probe details (linear, curved, etc), only specific item numbers for a specific vendor. | Those respective probes are needed. Equivalents are also fine.
580B60015 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | 7.a. Please confirm ICE catheter is needed or clinically necessary; if not, move to the value-add section. | It is needed.
580B60015 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | 8.a-h. Please confirm whether the fusion-targeted biopsy system, along with associated lines, is needed; if not, remove it or move it to the value-add section. | It is needed.
580B60015 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | G. Is there a trade in? | This item is an addition, no trade-ins.
580B60019 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | Request Item 1.a. Minimum Screen Size reduced from 20 in to 19 in. Request 6.i. Shear wave and strain-based elastography be removed. Under 8. Transducers can you verify this is the total transducers across all seven systems? | 1.a. Changed to 19in. 6.i. Moved to added value 8. Transducer list updated.
580B60019 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | Please remove request for specific vendor and ensure this spec is vendor neutral. | Specific ultrasound device requested, but no specific vendor requested
580B60019 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | 2.c. Please confirm need for 3D -freehand for use in the vascular lab. Request to have this moved to value add section or removed it not needed. | Moved to added value. FYSA: These ultrasound devices serve multiple departments and are not limited to vascular surgery.
580B60019 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | 2.d. Please confirm navigation software is needed. If not please remove or move to value add. This is typically a radiology procedure feature and not used in vascular surgery. | Moved to added value. FYSA: These ultrasound devices serve multiple departments and are not limited to vascular surgery.
580B60019 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | 2.d. Please confirm the image fusion software. If not needed, please remove or move to value-add. This is typically a radiology procedure feature and typically not used in vascular surgery. | Moved to added value. FYSA: These ultrasound devices serve multiple departments and are not limited to vascular surgery.
580B60019 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | 3.n Can the xPlane imaging feature be removed from the requirement as it appears to be brand-specific? | Moved to added value. FYSA: These ultrasound devices serve multiple departments and are not limited to vascular surgery.
580B60019 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | 6.i. Is shear-wave and/or strain-based elastography clinically needed for vascular surgery? If not, please remove or move to the value-add section. | Moved to added value. FYSA: These ultrasound devices serve multiple departments and are not limited to vascular surgery.
580B60019 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | 6.o. Please remove the phantom requirement or move it to the value-add section. This is usually a third-party purchase. | Removed.
580B60019 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | 7.g. Please confirm whether the mitral value navigator is clinically needed; if not, please remove it or move it to the value-add section. This is a cardiology feature, not vascular surgery. | Transducer list updated. FYSA: These ultrasound devices serve multiple departments and are not limited to vascular surgery.
580B60019 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | 7.h. Please confirm whether the cardiact 3D quantification is clinically needed; if not, please remove it or move it to the value-add section. This is a cardiology feature, not vascular surgery. | Removed. FYSA: These ultrasound devices serve multiple departments and are not limited to vascular surgery.
580B60019 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | 7.k Please confirm whether the 3D auto left atrial appendage n is clinically needed; if not, please remove it or move it to the value-add section. This is a cardiology feature, not a vascular surgery feature. | Removed.
580B60020 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | Missing specs for probes | List of the seven required transducers defined in section 7.
580B60020 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | Request item 1.m. Maximum equipment weight be increase from 250 to 265. Can you verify if 8.b. 3D-generated model of the prostate is required? | 1.m Increased 8.b. Moved to value add
580B60020 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | 2.b. Please confirm whether 3D freehand is needed for use at the CBOC. Request to have this moved to the value-add section or removed, as it is not needed. | 2.b. Moved to value add
580B60020 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | 2.c. Please confirm if navigation software is needed. If not, please remove or move to value-add. This is typically a radiology/IR department procedure feature and not used at the CBOC. | 2.c. Moved to value add
580B60020 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | 2.d. Please confirm if image fusion software is needed. If not, please remove or move to value-add. This is typically a radiology/IR department procedure feature and not used at the CBOC. | 2.d. Moved to value add
580B60020 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | 4.j. Please confirm if shear-wave and/or strain-based elastography clinically needed for the CBOC? I don't believe any of the Houston sites are doing shear-wave. If not, please remove or move to the value-add section. | 5.i. Required Specification
580B60020 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | 6.k. Please confirm whether the mitral value navigator is clinically needed; if not, please remove it or move it to the value-add section. This is a cardiology feature, not a radiology. | 6.k. Moved to value add
580B60020 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | 6.l. Please confirm whether the cardiact 3D quantification is clinically needed; if not, please remove it or move it to the value-add section. This is a cardiology feature, not radiology. | 6.l. Moved to value add
580B60020 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | 7. Please provide transducer required, spec is empty. | List of the 7 required transducers defined in section 7.
580B60020 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | 7. Is there a trade in? | No trade ins.
580B60022 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | Request Item 1.a. Minimum Screen Size reduced from 20 in to 19 in. Request 5.k. Shear wave and strain-based elastography be removed. | 1.a. Updated to 19 in. 5.k. Removed
580B60022 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | 5.k Please confirm whether shear-wave and/or strain-based elastography is clinically needed. This is a radiology feature, and I question whether it needs to be included, as the spec is for operative care and surgery. If not, please remove or move to the value-add section. | 5.k. Removed
580B60022 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | 5.o. Please remove the phantom requirement or move it to the value-add section. This is usually a third-party purchase. | 5.o. Removed
580B60022 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | g. Please confirm the manufacturer of all trade-ins; suspect these are Sonosites, but it says Phoenix Kiosk on some. | g. Updated
580B60023 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | Request item 1.m. Minimum equipment weight be increased to 265 lbs | 1.m Increased
580B60023 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | Please remove the request for a specific vendor and ensure this spec is vendor-neutral. These specs appear to be biased toward one vendor. | The specifications are vendor neutral.
580B60023 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | 2.b. Please confirm whether 3D freehand is needed for use at the CBOC. Request to have this moved to the value-add section or removed, as it is not needed. | 2.b. Moved to value add
580B60023 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | 2.c. Please confirm if navigation software is needed. If not, please remove or move to value-add. This is typically a radiology/IR department procedure feature and not used at the CBOC. | 2.c. Moved to value add
580B60023 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | 2.d. Please confirm if image fusion software is needed. If not, please remove or move to value-add. This is typically a radiology/IR department procedure feature and not used at the CBOC | 2.d. Moved to value add
580B60023 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | 3.n Can the xPlane imaging feature be removed from the requirement as it appears to be brand-specific? | 2.n. Moved to value add
580B60023 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | 5.j. Please confirm whether shear-wave and/or strain-based elastography is clinically needed for the CBOC? I don't believe any of the Houston sites are doing shear-wave. If not, please remove or move to the value-add section. | 5.j. Required Specification
580B60023 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | 7. Please provide transducer clarification; the section appears to be focused on one vendor without probe details (linear, curved, etc), only specific item numbers for a specific vendor. | 7. Added "or equivalent"
580B60023 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | 8.a. Can this Prostate Gland measurement be placed in the Value Add section? | 3. Moved to value add
580B60023 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | 2.b. Please confirm whether 3D freehand is needed for use at the CBOC. Request to have this moved to the value-add section or removed, as it is not needed. | 2.b. Moved to value add
580B60025 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | Please confirm the care area for this tech spec. Says anesthesia, but the request is only for a TEE probe and heavy cardiac features. | CONFIRMED, its for the Anesthesia Care Line
580B60025 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | 3.b. Please confirm whether 3D freehand is needed for use at the Anesthesia Request to have this moved to the value-add section or removed, as it is not needed. | Moved to Value-add section
580B60025 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | 3.c. Please remove navigation software, this is not applicable to anesthesia. | Removed
580B60025 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | 3.d. Please remove image fusion software, as this is not applicable in anesthesia but in radiology. | Removed
580B60025 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | 3.m. Can the X-Plane imaging feature be removed from the requirement as it appears to be brand-specific? | No it is needed by clinicians
580B60025 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | 6.i. Please confirm if shear-wave and/or strain-based elastography clinically needed. This feature is for radiology and not anesthesia. If not, please remove or move to the value-add section. | Removed
580B60025 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | 6.o. Please remove the phantom requirement or move it to the value-add section. This is usually a third-party purchase. | Moved to Added Value Section
580B60025 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | 7.g. Please confirm whether the mitral value navigator is clinically needed for anesthesia; if not, please remove it or move it to the value-add section. | Not required, has been removed.
580B60025 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | 7.h. Please confirm whether the cardiac 3D quantification is clinically needed for anesthesia; if not, please remove it or move it to the value-add section. | Is required by clinician
580B60025 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | 7.k. Please confirm whether the 3D auto left atrial appendage is clinically needed for anesethesia; if not, please remove it or move it to the value-add section | Moved to Value-add section
580B60025 | VISN 16: South Central VA Health Care Network | 580 HOUSTON TX | XR US | 8.a. Please confirm the transducers, as only a TEE probe is listed. Is this 2D or 3D? Are other probes needed for anesthesia? | The X8-2T, or equivalent probe is the only one needed. It is a 3D probe
583B60039 | VISN 10: VA Healthcare System of Ohio | 583 INDIANAPOLIS IN | XR US | Maximum weight per system for the 3 systems is stated at 24 lbs each, is the really the maximum weight for a system that is being requested for a structured heart program. Can this weight limit be actually be 240lbs each? | Yes, Please see revised Tech Specs and Compliance Matrix for changes.
583B60039 | VISN 10: VA Healthcare System of Ohio | 583 INDIANAPOLIS IN | XR US | Request 1.m. Maximum Weight be increased to 250 lbs. | Yes, Please see revised Tech Specs and Compliance Matrix for changes.
589B62002 | VISN 15: VA Heartland Network | 589 TOPEKA KS | XR PORT C ARM | Please adjust ninimum range of motion - panning to +/- 10 degrees for fair and open competition. | Requirement is for 18 degrees total, or +/- 9 degrees.
589B62002 | VISN 15: VA Heartland Network | 589 TOPEKA KS | XR PORT C ARM | Please adjust minimum vertical travel to 40 cm for fair and open competition | This is a clinical requirement and more than one vendor can meet it.
589B62002 | VISN 15: VA Heartland Network | 589 TOPEKA KS | XR PORT C ARM | Please adjust maximum weight for fair and open competition to 920 lbs | This is a clinical requirement and more than one vendor can meet it.
589B62002 | VISN 15: VA Heartland Network | 589 TOPEKA KS | XR PORT C ARM | Please remove UPS on board requirement for fair and open competition | Tech Spec Updated
589B62002 | VISN 15: VA Heartland Network | 589 TOPEKA KS | XR PORT C ARM | line 28 i. Can Minimum c-arm depth spec be changed to 26.5 for fair competition? | This is a clinical requirement and more than one vendor can meet it.
589B62002 | VISN 15: VA Heartland Network | 589 TOPEKA KS | XR PORT C ARM | line 54 u. Can this be changed to 20 seconds for fair competition, or moved to value add? | Tech Spec Updated
589B62002 | VISN 15: VA Heartland Network | 589 TOPEKA KS | XR PORT C ARM | line 61 c. This spec is requesting vascular features, but the matrix notes this system will be used for pain. Can you please confirm if device will be used for vascular procedures? | The device will be used for pain procedures, during which visualization of vasculature will be required.
589B63016 | VISN 15: VA Heartland Network | 589 COLUMBIA MO | XR US | Request 4.a. a keyboard on the display be required in lieu of the Physical keyboard? | Having a physical keyboard is a clinical requirement
589B63022 | VISN 15: VA Heartland Network | 589 COLUMBIA MO | XR US | 2.c. Please remove navigation software; this is not applicable for this care area. | Navigation software is a clinical requirement
589B63022 | VISN 15: VA Heartland Network | 589 COLUMBIA MO | XR US | 2.d. Please remove image fusion software; this is not applicable for this care area. | Image fusion is a clinical requirement
589B63022 | VISN 15: VA Heartland Network | 589 COLUMBIA MO | XR US | g. Are there any trade ins? | VA has no trade-in units to offer.
595B69013 | VISN 4: VA Stars & Stripes Healthcare Network | 595 LEBANON PA | XR RAD | 1.cb. Can you please change additional wireless detector to a 14x17 detector? 5.a. Please clarify if you would like to perform stitching at the wallstand and table.﻿ | 1 Yes, will change dectector size 5 Clarify in response what stitching is possible, both wallstand and table stitching is preferred
595B69013 | VISN 4: VA Stars & Stripes Healthcare Network | 595 LEBANON PA | XR RAD | 1.a.b. , 1.b.b. and 1.c.b. Can you change the minimum detector size to 14 x 17 as 15 x 15 is not a common detector size and this will increase competition? | 1 Yes, will change dectector size
595B69013 | VISN 4: VA Stars & Stripes Healthcare Network | 595 LEBANON PA | XR RAD | 2.c. Can you change the maximum pixel size to 140 to increase competition? | No. Equipment weight cannot be increased. It is necessary to support mobility of the ultrasound as well as minimize strain on end users.
598B60018 | VISN 16: South Central VA Health Care Network | 598 LITTLE ROCK AR | XR MAMMO | Line 21A, 22B The specifications requires 20-39kV(2D) and 20-49kV(3D); will systems with alternative manufacturer designed operating ranges and equivalent clinical performance be considered acceptable? | Yes
598B60018 | VISN 16: South Central VA Health Care Network | 598 LITTLE ROCK AR | XR MAMMO | Line 32L The specifications requires 0.1MM; will systems with equivalent image quality and lesion visualization be considered acceptable? | Yes
598B60018 | VISN 16: South Central VA Health Care Network | 598 LITTLE ROCK AR | XR MAMMO | Line 33M Would a 65 cm SID design be considered a clinically equivalent or substantially compliant alternative to the specified 70cm SID | Please submit according to your best understanding. The best value determination will make considerations where feasible.
598B60018 | VISN 16: South Central VA Health Care Network | 598 LITTLE ROCK AR | XR MAMMO | Line 48H, 49I Please clarify whether FDA-cleared third party CAD and AI tomosynthesis detection solutions integrated with the mammography workflow satisfy the CAD and AI detection requirements | Please submit according to your best understanding. The best value determination will make considerations where feasible.
598B60018 | VISN 16: South Central VA Health Care Network | 598 LITTLE ROCK AR | XR MAMMO | 4a: Biopsy Paddle-There is no biopsy system listed on specs, are you requesting a needle loc kit? | Not requested at this time.
598B60018 | VISN 16: South Central VA Health Care Network | 598 LITTLE ROCK AR | XR MAMMO | 1.Unit physical specs, i: Please confirm that you are requesting high resolution imaging. | Yes
598B60018 | VISN 16: South Central VA Health Care Network | 598 LITTLE ROCK AR | XR MAMMO | 3h. Computer-aided detection: Is this 2D CAD? 2D & 3D CAD? Or 3D CAD? | 2D & 3D
598B60018 | VISN 16: South Central VA Health Care Network | 598 LITTLE ROCK AR | XR MAMMO | 1. Do you need a biopsy chair quoted? 2. What vacuum assisted device will you be using for biopsy? 3. Will the site need needle localization paddles and the associated tools? | 1. no 2. Group has BARD Magnum unit 3. no
598B60020 | VISN 16: South Central VA Health Care Network | 598 LITTLE ROCK AR | XR MAMMO | Line 21A, 22B The specifications requires 20-39kV(2D) and 20-49kV(3D); will systems with alternative manufacturer designed operating ranges and equivalent clinical performance be considered acceptable? | Yes
598B60020 | VISN 16: South Central VA Health Care Network | 598 LITTLE ROCK AR | XR MAMMO | Line 32L The specifications requires 0.1MM; will systems with equivalent image quality and lesion visualization be considered acceptable? | Yes
598B60020 | VISN 16: South Central VA Health Care Network | 598 LITTLE ROCK AR | XR MAMMO | Line 33M Would a 65 cm SID design be considered a clinically equivalent or substantially compliant alternative to the specified 70cm SID | Yes
598B60020 | VISN 16: South Central VA Health Care Network | 598 LITTLE ROCK AR | XR MAMMO | Line 48H, 49I Please clarify whether FDA-cleared third party CAD and AI tomosynthesis detection solutions integrated with the mammography workflow satisfy the CAD and AI detection requirements | Yes
598B60020 | VISN 16: South Central VA Health Care Network | 598 LITTLE ROCK AR | XR MAMMO | 4a: Biopsy Paddle-There is no biopsy system listed on specs, are you requesting a needle loc kit? | Not requested at this time.
598B60020 | VISN 16: South Central VA Health Care Network | 598 LITTLE ROCK AR | XR MAMMO | Please provide the correct system trade-in SN. It will start with 810. | Mammo 1 (PO#598B60018): MX1157252/EE117092, SN: 81108155771, 81208155589 & 81008155337 Console: SN: 81008155337 Mammo 2 (PO#598B60020): MX1133707/EE122844 SN: 8100217793, SN:81003177983, SN: 81102178657 Console: SN: 81412167681
598B60020 | VISN 16: South Central VA Health Care Network | 598 LITTLE ROCK AR | XR MAMMO | 1.Unit physical specs, i: Please confirm that you are requesting high resolution imaging. | Yes
598B60020 | VISN 16: South Central VA Health Care Network | 598 LITTLE ROCK AR | XR MAMMO | 3h. Computer-aided detection: Is this 2D CAD? 2D & 3D CAD? Or 3D CAD? | 2D & 3D
598B60020 | VISN 16: South Central VA Health Care Network | 598 LITTLE ROCK AR | XR MAMMO | 1. Do you need a biopsy chair quoted? 2. What vacuum assisted device will you be using for biopsy? 3. Will the site need needle localization paddles and the associated tools? | 1. no 2. Group has BARD Magnum unit 3. no
648B63006 | VISN 20: VA Northwest Network | 648 PORTLAND OR | XR US | Can you please confirm functionality of the system as the description details robotic kidney and open kidney surgeries but does not include robotic analysis package or applicable transducers. Image Fusion Software. Please confirm what procedures will be performed utilizing Fusion. If Fusion is required, will it be needed for TP, TR, or both? If Fusion is required, will it be required to be freehand or stabilized? Can 1. Clinical Training d. Technologists who complete the clinical applications training shall receive continuing education credits (CMEs) be added value? Is the trade in operational? | 1. Confirmed, additional application packages may be offered as optionalImage Fusion Software. 2. Please confirm what procedures will be performed utilizing Fusion Prostate biopsies - TRUS biopsies 3. If Fusion is required, will it be needed for TP, TR, or both? TR 4. If Fusion is required, will it be required to be freehand or stabilized? Freehand. 5. This will not be moved to added value 6. Trade in is operational
621B60002 | VISN 9: VA Mid South Healthcare Network | 621 MOUNTAIN HOME TN | XR US | 1.j. Please change minimum number of transducer ports to “4”, and remove “parking ports”. | A minimum of 4 active transducer ports is required, including 2 parking ports.
626B64028 | VISN 9: VA Mid South Healthcare Network | 626 CHATTANOOGA TN | XR US | Request 1.k. Minimum Battery life be reduced to 30 min. | 45 minute battery life minimum is required for ease of use of clinical staff, movement of ultrasound between use, and in the event of power outage.
626B64028 | VISN 9: VA Mid South Healthcare Network | 626 CHATTANOOGA TN | XR US | Section 6, Analysis Packages. B. Abdominal: Would they like to include our Abdominal C5-1 probe since they are adding our Abdominal clinical option file? | No abdominal probe needed.
626B64028 | VISN 9: VA Mid South Healthcare Network | 626 CHATTANOOGA TN | XR US | Section 6, Analysis Packages. C. Vascular: Would they like our Vascular Tools - 1.) MicroFlow Imaging and 2.) Flow Viewer? | No, vascular tools are not required.
626B64028 | VISN 9: VA Mid South Healthcare Network | 626 CHATTANOOGA TN | XR US | Section 7, Transducers. Since Vascular option is requested, Would they like to include our L12-3 Ergo Linear probe and L15-7io Linear probe? | No, the L12-3 and L15-7io probes are not required.
626B64028 | VISN 9: VA Mid South Healthcare Network | 626 CHATTANOOGA TN | XR US | Shear Wave and strain based elastrography Can you please clarify the intended cardiovascular clinical application for shear wave and strain-based elastography within this requirement? Is this referring to cardiac strain imaging, or is tissue elastography specifically required? | This is referring to strain imaging.
632B61006 | VISN 2: VA Healthcare Network Upstate New York | 632 NORTHPORT NY | XR US | Request Item 1.a. Minimum Screen Size reduced from 20 in to 19 in. | Reduced to 19 in
632B61006 | VISN 2: VA Healthcare Network Upstate New York | 632 NORTHPORT NY | XR US | 1.h. Can the minimum battery life be reduced to 80 mins to increase competition? | Reduced to 80 mins
635B60002 | VISN 19: VA Rocky Mountain Network | 635 OKLAHOMA CITY OK | XR US | Schedule Line Items: There are two line item with the same obligation numbers, only one Matrix and Technical Specifications were sent for one ultrasound. Is there another ultrasound for Oklahoma or is this a misprint ? | This was a misprint; there should just be 1 line item.
635B60002 | VISN 19: VA Rocky Mountain Network | 635 OKLAHOMA CITY OK | XR US | 3.b. Please confirm if navigation software is needed. If not, please remove or move to value-add. | Navigation software is not needed.
635B60002 | VISN 19: VA Rocky Mountain Network | 635 OKLAHOMA CITY OK | XR US | 6.d. Please confirm if shear-wave and/or strain-based elastography is clinically needed for the OR/ICU. This feature is found more often in radiology. If not, please remove or move to the value-add section. | Shear-wave and/or strain-based elastography moved to value-add.
635B60003 | VISN 19: VA Rocky Mountain Network | 635 OKLAHOMA CITY OK | XR PORT C ARM | Please adjust minimum vertical travel to 40 cm for fair and open competition | This can be adjusted – minimum vertical travel changed to 40cm from 45cm.
635B60003 | VISN 19: VA Rocky Mountain Network | 635 OKLAHOMA CITY OK | XR PORT C ARM | Please adjust maximum height requirement to 79" for fair and open competition | This can be adjusted – maximum height increased to 80” from 75”.
635B60003 | VISN 19: VA Rocky Mountain Network | 635 OKLAHOMA CITY OK | XR PORT C ARM | Please adjust maximum weight for fair and open competition to 920 lbs | Maximum weight increased to 950 lbs
637B69003 | VISN 6: VA Mid-Atlantic Health Care Network | 637 ASHEVILLE NC | NM PET CT | Can the following information be provided about the trade-in: • Date of Manufacture • Serial Number • Crane needed? • System in good working condition? • Last Tube exchange • Scan seconds on CT • Tube in working condition? • Software level | Date of Manufacture: July 2015 Serial Number: 25202 Crane needed? No System in good working condition? Yes Last Tube exchange: 14 Aug 2020 Scan seconds on CT: 1d 13h 17m 18s (total on tube) Tube in working condition? Yes Software level: 1.003
637B69003 | VISN 6: VA Mid-Atlantic Health Care Network | 637 ASHEVILLE NC | NM PET CT | 1h. Please convert 1.h to minimum crystal volume at 10,880,000 mm³ (34K x (4mm x 4mm x20mm)). Crystal volume or NEMA NU 2-2018 sensitivity is a more objective measure of detector material available for photon detection versus crystal count. 1i. Please update to detector design shall provide clinical spatial resolution of 3.7mm achieved by 4mm × 4mm × 20mm crystals, as demonstrated through published performance specifications and standard clinical reconstruction methods. 1j. Please change to minimum effective sensitivity of 128 cps/kBq, or equivalent lesion detectability, image quality, and dose efficiency demonstrated through published manufacturer specifications and accepted industry performance standards. 1k. Time of Flight localization is typically measured in cms. Is this a duplicate for technical requirement? 1j. and if so, can you please change to minimum effective sensitivity of 128 cps/kBq, or equivalent lesion detectability, image quality, and dose efficiency demonstrated through published manufacturer specifications and accepted industry performance standards. 1l. Can this specification be updated to reflect peak NECR (Noise Equivalent Count Rate), which is a more comprehensive measure of system count-rate performance and clinical imaging capability, with the timing window serving as one contributing factor? 2l. Would the customer consider a Digital BGO-based' PET/CT System?" | 1h - Remaining as Min Number of Crystals 1i - Changed to "Max Crystal Size" - 4mm x 4mm x 20mm 1j - Remaining as TOF Resolution 1k - Deleted 1l - Remaining as TOF Coincidence Window 2l - No
648B63007 | VISN 20: VA Northwest Network | 648 PORTLAND OR | XR US | Request item 1.k. Maximum weight be increased to 165 lbs | No
668B63003 | VISN 20: VA Northwest Network | 668 SPOKANE WA | XR US | Request item 1.a. Minimum screen size reduced to 15.6 in. | Yes
612B69010 | VISN 21: VA Sierra Pacific Network | 612 MARINEZ CA | XR US | Request item 1.k. Maximum weight be increased to 145 lbs Request item 1.i. Minimum Battery life be revised to 120 min. Request 5.p. Voice Activation be removed or moved to added value as it is restrictive. | 1.k Maximum weight increased to increase competition. 1.i. Minimum battery life will not be decreased 5.p. Voice activation moved to value added to increase competition.
640B64053 | VISN 21: VA Sierra Pacific Network | 640 PALO ALTO CA | XR US | Request item 1.g. Maximum weight be increased to 145 lbs Request 5.g. Voice Activation be removed or moved to added value as it is restrictive. | 1.g increased to 145. 5.g Voice Activatio moved to added value
640B64054 | VISN 21: VA Sierra Pacific Network | 640 PALO ALTO CA | XR US | Request item 1.g. Maximum weight be increased to 145 lbs Request 5.g. Voice Activation be removed or moved to added value as it is restrictive. | 1.g increased to 145. 5.g Voice Activatio moved to added value
640B64056 | VISN 21: VA Sierra Pacific Network | 640 PALO ALTO CA | XR US | Request item 1.g. Maximum weight be increased to 145 lbs | Increased to 145
640B64056 | VISN 21: VA Sierra Pacific Network | 640 PALO ALTO CA | XR US | 2.b. Please confirm whether 3D freehand is needed for use at the hand clinic. If not, please remove or move to the value-add section. | 3D is a requirement
640B64056 | VISN 21: VA Sierra Pacific Network | 640 PALO ALTO CA | XR US | 2.c. Please remove navigation software; this is not applicable for this care area. | Removed
640B64056 | VISN 21: VA Sierra Pacific Network | 640 PALO ALTO CA | XR US | 7a. Can the frequency of the linear be reduced to 22Mhz or lower to allow for more competition than just one vendor? | Reduced to 22Mhz
640B64057 | VISN 21: VA Sierra Pacific Network | 640 PALO ALTO CA | XR PORT C ARM | line 31 l. Can the maximum system height be adjusted to 70.9"? | No. Limited clinical space requires compact system.
640B64057 | VISN 21: VA Sierra Pacific Network | 640 PALO ALTO CA | XR PORT C ARM | line 32 m. Can the maximum system footprint be adjusted to 29' x 39" | No. Limited clinical space requires compact system.
640B64057 | VISN 21: VA Sierra Pacific Network | 640 PALO ALTO CA | XR PORT C ARM | line 33 n. Can the maximum system weight be changed to 485lbs? | No. Multiple vendors can meet this requirements.
640B64057 | VISN 21: VA Sierra Pacific Network | 640 PALO ALTO CA | XR PORT C ARM | line 48 n. Can the remote control option for entire unit be moved to value add for fair competition? | No. Multiple vendors can meet this requirements.
640B64057 | VISN 21: VA Sierra Pacific Network | 640 PALO ALTO CA | XR PORT C ARM | line 23 d. Can the minimum range of motion - orbital rotation be changed to 120 degrees? | No. Needed for Podiatry and Hand surgeries. Decreasing rotation decreases surgeries system can be used for.
640B64057 | VISN 21: VA Sierra Pacific Network | 640 PALO ALTO CA | XR PORT C ARM | line 24 e. can the minimum range f motion pivotol rotation be changed to 380 degrees? | No. Needed for Podiatry and Hand surgeries. Decreasing rotation decreases surgeries system can be used for.
640B64057 | VISN 21: VA Sierra Pacific Network | 640 PALO ALTO CA | XR PORT C ARM | line 27 h. can the minimum c-arm arc depth be changed to 18"? | No. Multiple vendors can meet this requirements.
640B64058 | VISN 21: VA Sierra Pacific Network | 640 PALO ALTO CA | XR PORT C ARM | line 22 c. can the minimum detector size spec be changed to 15 cm x 15 cm for fair competition? | No. Needed for specific extremity surgeries. Decreasing detector size eliminates ability to use system for these surgeries.
640B64058 | VISN 21: VA Sierra Pacific Network | 640 PALO ALTO CA | XR PORT C ARM | line 23 d. Can the minimum range of motion - orbital rotation be changed to 120 degrees? | No. Needed for Podiatry and Hand surgeries. Decreasing rotation decreases surgeries system can be used for.
640B64058 | VISN 21: VA Sierra Pacific Network | 640 PALO ALTO CA | XR PORT C ARM | line 24 e. can the minimum range f motion pivotol rotation be changed to 380 degrees? | No. Needed for Podiatry and Hand surgeries. Decreasing rotation decreases surgeries system can be used for.
640B64058 | VISN 21: VA Sierra Pacific Network | 640 PALO ALTO CA | XR PORT C ARM | line 27 h. can the minimum c-arm arc depth be changed to 18"? | No. Multiple vendors can meet this requirements.
640B64058 | VISN 21: VA Sierra Pacific Network | 640 PALO ALTO CA | XR PORT C ARM | line 28 i. Can the minimum monitor size be changed to 27"? | No.
640B64058 | VISN 21: VA Sierra Pacific Network | 640 PALO ALTO CA | XR PORT C ARM | line 31 l. Can the maximum system height be adjusted to 70.9"? | No. Limited clinical space requires compact system.
642B69014 | VISN 4: VA Stars & Stripes Healthcare Network | 642 PHILADELPHIA PA | XR US | Request item 1.j. Maximum equipment weight to be increased to 265 lbs | No. Equipment weight cannot be increased. It is necessary to support mobility of the ultrasound as well as minimize strain on end users.
640B64058 | VISN 21: VA Sierra Pacific Network | 640 PALO ALTO CA | XR PORT C ARM | line 32 m. Can the maximum system footprint be adjusted to 29' x 39" | No. Limited clinical space requires compact system.
640B64058 | VISN 21: VA Sierra Pacific Network | 640 PALO ALTO CA | XR PORT C ARM | line 33 n. Can the maximum system weight be changed to 485lbs? | No. Multiple vendors can meet this requirements.
640B64058 | VISN 21: VA Sierra Pacific Network | 640 PALO ALTO CA | XR PORT C ARM | line 48 n. Can the remote control option for entire unit be moved to value add for fair competition? | No. Multiple vendors can meet this requirements.
662B60014 | VISN 21: VA Sierra Pacific Network | 662 SAN FRANCISCO CA | XR US | Request 4.a. a keyboard on the display be required in lieu of the Physical keyboard? | Yes
662B60014 | VISN 21: VA Sierra Pacific Network | 662 SAN FRANCISCO CA | XR US | QTY. 5-1 .Abdomen, Cardiac, Focused Cardiac, Lung, OB, Orbital, TCD. Can we please confirm is this for s5-1 or x5-1 transducer. Also do they need stress or strain? | s5-1 and No
662B60015 | VISN 21: VA Sierra Pacific Network | 662 SAN FRANCISCO CA | XR US | Column A does not include this Equipment PO - 662B60025 in order to pre populate the VISN, Station and Equipment type. Missing specs benchmark | Tech specs have been provided to the NAC
662B60015 | VISN 21: VA Sierra Pacific Network | 662 SAN FRANCISCO CA | XR US | Column A does not include this Equipment PO - 662B60015 in order to pre populate the VISN, Station and Equipment type. Missing specs benchmark | Tech specs have been provided to the NAC
662B60025 | VISN 21: VA Sierra Pacific Network | 662 SAN FRANCISCO CA | XR US | Benchmark Missing | Tech specs have been provided to the NAC
662B60025 | VISN 21: VA Sierra Pacific Network | 662 SAN FRANCISCO CA | XR US | please supply the technical requirements and matrix | Tech specs have been provided to the NAC
662B60031 | VISN 21: VA Sierra Pacific Network | 662 SAN FRANCISCO CA | XR PORT C ARM | line 15 Delivery address can you please confirm if there is a suite number for this address, appears to be a large building with multpile facilities | Please use the delievery address on the specs. The VA will use the same warehouse as everyone else in the building
662B60031 | VISN 21: VA Sierra Pacific Network | 662 SAN FRANCISCO CA | XR PORT C ARM | line 53 t. Can the rotating flat detector be moved to value add or can the ability to rotate the image at the tube head or display monitor be included as part of the spec? | Please see updated tech specs
644B68023 | VISN 22: VA Desert Pacific Healthcare Network | 644 PHOENIX AZ | XR PORT C ARM | Please adjust orbital rotation to -120/ +45 for fair and open competition | Orbital rotation adjusted
659B69001 | VISN 6: VA Mid-Atlantic Health Care Network | 659 SALISBURY NC | XR US | Can the maximum product weight be expanded to 200lbs to increase competition? Can you please confirm functionality of the system as there is not description and no applications listed. Can 1. Clinical Training d. Technologists who complete the clinical applications training shall receive continuing education credits (CMEs) be added value? | Weight will remain at 150 lbs, multiple vendors can meet requirement This system is supposed to be a standard Point-of-Care ultrasound system that is intended to held with needle insertion in intra operative use cases. No; clinical training d. will not be moved to added value.
644B68023 | VISN 22: VA Desert Pacific Healthcare Network | 644 PHOENIX AZ | XR PORT C ARM | Please adjust minimum c-arm depth to 27" to allow for fair and open competition | C-arm depth is not a spec listed
644B68023 | VISN 22: VA Desert Pacific Healthcare Network | 644 PHOENIX AZ | XR PORT C ARM | Please adjust maximum weight for fair and open competition to 920 lbs | Maximum weight adjusted to 650
644B68023 | VISN 22: VA Desert Pacific Healthcare Network | 644 PHOENIX AZ | XR PORT C ARM | line 26 g. Can the minimum free space be changed to 79" for fair competion? | Free space adjusted to 79"
644B68023 | VISN 22: VA Desert Pacific Healthcare Network | 644 PHOENIX AZ | XR PORT C ARM | line 31 l. Can the maximum weight be changed to 650 for fair competition? | Maximum weight adjusted to 650
644B68024 | VISN 22: VA Desert Pacific Healthcare Network | 644 PHOENIX AZ | XR PORT C ARM | Please adjust minimum requirement for orbital rotation to 165 degrees for fair and open competition | Orbital rotation adjusted
644B68024 | VISN 22: VA Desert Pacific Healthcare Network | 644 PHOENIX AZ | XR PORT C ARM | Please adjust minimum vertical travel to 40 cm for fair and open competition | Vertical travel adjusted
644B68025 | VISN 22: VA Desert Pacific Healthcare Network | 644 PHOENIX AZ | XR PORT C ARM | Please adjust minimum requirement for orbital rotation to 165 degrees for fair and open competition | Orbital rotation adjusted
644B68025 | VISN 22: VA Desert Pacific Healthcare Network | 644 PHOENIX AZ | XR PORT C ARM | Please adjust minimum vertical travel to 40 cm for fair and open competition | Vertical travel adjusted
644B68025 | VISN 22: VA Desert Pacific Healthcare Network | 644 PHOENIX AZ | XR PORT C ARM | Please remove UPS on board requirement for fair and open competition | UPS on board is needed.
664B60058 | VISN 22: VA Desert Pacific Healthcare Network | 664 SAN DIEGO CA | XR US | Request Item 1.a. Minimum Screen Size reduced from 20 in to 19 in. Request 4.a. a keyboard on the display be required in lieu of the Physical keyboard? | 1a. Concur with change to 19 in. specifications updated 4a. Physical keyboard required.
671B60042 | VISN 17: VA Heart of Texas Health Care Network | 671 SAN ANTONIO TX | XR PORT C ARM | Please clarify if a flat panel or image intesifier is required | The flat panel is required
671B60042 | VISN 17: VA Heart of Texas Health Care Network | 671 SAN ANTONIO TX | XR PORT C ARM | Please clarify if 32" is the required monitor size per monitor or for one single monitor | 32 inch is the required monitor size
671B60042 | VISN 17: VA Heart of Texas Health Care Network | 671 SAN ANTONIO TX | XR PORT C ARM | Please clarify minimum horizontal travel, no c-arm meets 200 cm or 78 inches | Requirement has been changed to 20cm
671B60042 | VISN 17: VA Heart of Texas Health Care Network | 671 SAN ANTONIO TX | XR PORT C ARM | Please adjust maximum weight for fair and open competition to 920 lbs | The requirement has been changed to 920lbs
671B60042 | VISN 17: VA Heart of Texas Health Care Network | 671 SAN ANTONIO TX | XR PORT C ARM | line 24 e. Can you please confirm if the maximum spatial resolution is meant to be in pixels? Please clarify the spec you are looking for. | The requirement has been changed to 2.8lp/mm
671B60042 | VISN 17: VA Heart of Texas Health Care Network | 671 SAN ANTONIO TX | XR PORT C ARM | line 28 i. Can you please confirm if horizontal travel should be in mm or cm? For fair competition can this be changed to 20cm? | Requirement has been changed to 20cm
671B60042 | VISN 17: VA Heart of Texas Health Care Network | 671 SAN ANTONIO TX | XR PORT C ARM | line 29 j. Can you please confirm if the vertical travel should be in mm or cm, can the vertical travel spec be changed to either 460mm or 46cm? | The requirement has been changed to 46cm
671B60042 | VISN 17: VA Heart of Texas Health Care Network | 671 SAN ANTONIO TX | XR PORT C ARM | line 32 m. Can the maximum system dimension be changed to 83.1" for fair competition? | The requirement has been changed to 84in
671B60042 | VISN 17: VA Heart of Texas Health Care Network | 671 SAN ANTONIO TX | XR PORT C ARM | line 33 n. Can the max weight be adjusted to 735 lbs for fair competition? | The requirement has been changed to 920lbs
672FC6012 | VISN 8: VA Sunshine Healthcare Network | 672 SAN JUAN PR | XR HYBRID | 1. Is this a sole source or an actual competitive bid. a. Offeror shows the Philips Sales rep. b. Manufacture and Model are already filled in as Philips c. Quote number already with a Philips quote number d. Multiple specifications using Philips nomenclature. e. Multiple specifications are Philips-only technologies. 2. UPS for 180 minutes-there is no UPS that I am aware of that can supply power for a full system for 180 minutes. The standard is 7-12 minutes. 3. Laser Guiding system—what’s the clinical purpose? 4. Line 3 Computed Tomography Specifications- section is blank. Is this a request for a traditional hybrid room or is it a Angio/CT combination hybrid lab? 5. Philips IVUS limits either Philips providing Volcano or Philips providing a quote to another vendor at a much higher price. Is there a way to increase competition at best value for the government? | 1. Competitive Bid- All referneces to a vendor have been removed. 2. 12 min 3. To display red laser lines visible displayed on patient, to correctly align equipment above patient. 4. This is a traditional hybrid room. We require a C-arm that can perform helical scans and 3D reconstruction. 5. Provider requires intravascular ultrasound.
675B67006 | VISN 8: VA Sunshine Healthcare Network | 675 ORLANDO FL | XR MRI | • Minimum field of view lists 70cm, can you explain what is meant by a 70cm minimum field of view • RF channels, 32 channels is listed as the minimum is this within a single field of view? • Should additional channel availability be shown as options? • Elastography typically includes hardware and software. Is it desired that there are two additional hardware driver sets for elastography • There are two requirements for UPS, a full UPS covers the MR system and computer workstation to confirm a second listed on line 69 is not required ? • What is the VA’s expectation of FUNCTIONAL head coil beyond the standard 20 channel or tiltable head or neurovascular coils? | 1. 70cm Open Bore Design 2. Min 64 RF channels 3.Yes 4.Elastography is requested #3 5.UPS is requested for complete MRI system 6.Min 20 channel head coil is requested with addition a send and recieve head coild option for implant patients
675B67006 | VISN 8: VA Sunshine Healthcare Network | 675 ORLANDO FL | XR MRI | 3.l. please change shielding to Active Shielding to comply with current industry standards 5.a. please confirm whether 2 torso coils are required or if 1 torso coil with 60cm of coverage is adequate | 1. Active shielding updated 2. Min 40cm and 60cm torso coil coverage options are requested
675B67007 | VISN 8: VA Sunshine Healthcare Network | 675 ORLANDO FL | XR IR | Philips IVUS limits either Philips providing Volcano or Philips providing a quote to another vendor at a much higher price. Is there a way to increase competition at best value for the government? | IVUS is the clinical requirement for IR.
675B67007 | VISN 8: VA Sunshine Healthcare Network | 675 ORLANDO FL | XR IR | 5.a.,d.,e.,j.,m.: please move to Value Adds section to be more vendor inclusive. 5.n. Please remove IVUS as this is a vendor-specific platform. 4.z. Please move to Value Adds section. 4.b. Ceiling mounted additional specification be removed or added to the Value Adds sections. 3.b. Please change minimum size monitor to 55in. 1.a. Please change minimum detector size to 30x30. | 1a: Detector size can't be smaller than 40x30. 3b: Change made 4b, z: These are necessary specification that's used during procedures. 5.a.,d.,e.,j.,m.: These software requirements are necessary for the functionality of device according to the procedures that are conducted on a daily basis. 5n: The ultrasound is a necessary component that's used with the main imaging during procedures.
675B67012 | VISN 8: VA Sunshine Healthcare Network | 675 ORLANDO FL | XR US | Request item 1.a. and 1.b. Minimum screen size be revised to 19 in. Can you verify that 2.c Navigation and 2.d. Image Fusion will be used in the ED? Can you verify that Elastography will be in used in the ED? | Navigation is needed, but Image Fusion will not be used in the ED. No elastography will be used in the ED. Minimum screen size was updated to 19 inches.
675B67014 | VISN 8: VA Sunshine Healthcare Network | 675 ORLANDO FL | XR US | Request item 1.a. and 1.b. Minimum screen size be revised to 19 in. Request item 8.b, and 8.c. be removed as not typically done in Pain Management. | Minimum screen size was updated to 19 inches. 8.b and 8.c will not be removed as these software services were requested by the clinical users for future expansion of services.
675B67015 | VISN 8: VA Sunshine Healthcare Network | 675 ORLANDO FL | XR US | Request item 1.a. and 1.b. Minimum screen size be revised to 19 in. Request item 8.b, and 8.c. be removed as not typically done in Inpatient Care. | Minimum screen size was updated to 19 inches. 8.b and 8.c will not be removed as these software services were requested by the clinical users for future expansion of services.
675B67016 | VISN 8: VA Sunshine Healthcare Network | 675 ORLANDO FL | XR PORT C ARM | Please adjust orbital rotation to -120/ +45 for fair and open competition | This is a necessary specification for system functionality.
675B67016 | VISN 8: VA Sunshine Healthcare Network | 675 ORLANDO FL | XR PORT C ARM | Please adjust minimum vertical travel to 40 cm for fair and open competition | This is a necessary specification for system functionality.
675B67016 | VISN 8: VA Sunshine Healthcare Network | 675 ORLANDO FL | XR PORT C ARM | Please adjust minimum c-arm depth to 27" to allow for fair and open competition | This is a necessary specification for system functionality.
675B67016 | VISN 8: VA Sunshine Healthcare Network | 675 ORLANDO FL | XR PORT C ARM | Please adjust maximum height requirement to 79" for fair and open competition | This is a necessary specification for system functionality.
675B67016 | VISN 8: VA Sunshine Healthcare Network | 675 ORLANDO FL | XR PORT C ARM | Please remove UPS on board requirement for fair and open competition | This is a necessary specification for system functionality.
675B67017 | VISN 8: VA Sunshine Healthcare Network | 675 ORLANDO FL | XR MRI | • Minimum field of view lists 70cm, can you explain what is meant by a 70cm minimum field of view • RF channels, 32 channels is listed as the minimum is this within a single field of view? • Should additional channel availability be shown as options? • Elastography typically includes hardware and software. Is it desired that there are two additional hardware driver sets for elastography • There are two requirements for UPS, a full UPS covers the MR system and computer workstation to confirm a second listed on line 69 is not required ? • What is the VA’s expectation of FUNCTIONAL head coil beyond the standard 20 channel or tiltable head or neurovascular coils? | 1. 70cm open bore design 2. Min 64 RF channels 3. YES 4. Elastography is requested on one of the two scanners requested. #3 only 5.UPS is requested for complete MRI system 6. Min 20 channel head coil reqested with additional send and receive head coil option for implant patient imaging.
675B67017 | VISN 8: VA Sunshine Healthcare Network | 675 ORLANDO FL | XR MRI | 3.l. Please change shielding to Active Shielding to comply with current industry standards. 5.a. Please confirm whether two torso coils are required or if one torso coil with 60cm of coverage is adequate. | 1. Active shielding updated 2. Min 40 and 60cm torso coil coverage options are requested
675B67017 | VISN 8: VA Sunshine Healthcare Network | 675 ORLANDO FL | XR MRI | Is this new 3.0T MR planned installation inside the current VA building? Will this new 3.0T install in a modular building? Will this new 3.0T install in a surrounding building already on the property? | 1. No, this will be a new building that will begin construction soon. 2. This will not be a modular building. 3. No, this will be a new building that will begin construction soon.
675B67017 | VISN 8: VA Sunshine Healthcare Network | 675 ORLANDO FL | XR MRI | Is this new 3.0T MR planned installation inside the current VA building? Will this new 3.0T install in a modular building? Will this new 3.0T install in a surrounding building already on the property? Compressed Sense speed is a line item. Does the VA not want to implement AI tools? What level of multi-transmit is expected for this 3.0T? Is multi-transmit supported by 2-RF transmitters, the baseline for this 3.0T? | 1. The two new MRI units will be housed in a new building that is scheduled to begin construction soon. 2. This will not be a modular building. 3. No, this will be a new building that will begin construction soon. 4. Compressed sense and AI technology is requested 5. Send and recieve brain coil is requested for implant imaging 6.No
678B61031 | VISN 22: VA Desert Pacific Healthcare Network | 678 TUCSON AZ | XR CT | 1.b. Please change aperture to 80 cm to be vendor inclusive. 1.c. Would 40 mm or 80 mm detector width be acceptable? 1.e. Please change minimum reconstructed slice to 0.4 mm to be more vendor inclusive 1.j. Please change the depth dimension to 57 to make this more vendor inclusive. | 1b. Concur with change to minimum 80 cm 1c. Minimum detector width 40mm is acceptable 1e. No, there are at least 2 vendors that meet requirment 1j. No, there are at least 2 vendors that meet requirement
678B61032 | VISN 22: VA Desert Pacific Healthcare Network | 678 TUCSON AZ | XR MRI | 1.f. Minimum number of channels 204. Please change to 64 channels for full competition and to reflect the industry standard. Maximum number of channels per Field of View without table movement each generate an independent partial image, not simply the number of elements that can be connected. 2.c. Minimum patient table load capacity [lbs 550]: Please change to 500 lbs for full competition. 4.c. Minimum number of processing/reading workstations 3: please clarify if an acquisition workstation and post processing workstation are required. Some acquisition workstations only use one wide-screen monitor. 4.f. Workstation UPS that maintains full functionality for a minimum of 60 minutes during a power outage: Please change to 7 minutes as we are not aware of a UPS that provides 60 minutes of functionality. 7.o. High Intensity Focused Ultrasound: Please clarify if the system simply needs to provide compatibility with third party HiFU devices. | 1f. No, more than 1 vendor capable of meeting requirement. 2.c. Concur with decrease to min. 500 lbs 4.c.1 acquisition station with 1 24in monitor will suffice 4.f 7 minutes is acceptable as the full system UPS provides redundancy. 7.o. Requires the HIFU system to be included with delivery.
688B61003 | VISN 5: VA Capitol Health Care Network | 688 WASHINGTON DC | XR US | Can the minimum monitor size be updated to 19 inches to increase competition? Can the maximum product weight be expanded to 200lbs to increase competition? Image Fusion Software. Please confirm what procedures will be performed utilizing Fusion. If Fusion is required, will it be needed for TP, TR, or both? If Fusion is required, will it be required to be freehand or stabilized? | Minimum monitor size updated to 19 inches in tech specs. Maximum equipment weight was already 200 lbs. No changes required. Fusion will be used for freehand TP
691B60052 | VISN 22: VA Desert Pacific Healthcare Network | 691 LOS ANGELES CA | XR PORT C ARM | Please clarify if a flat panel or image intesifier is required | Flat panel detector is required.
691B60052 | VISN 22: VA Desert Pacific Healthcare Network | 691 LOS ANGELES CA | XR PORT C ARM | Please adjust minimum horizontal travel to 20 cm for fair and open competition | Adjusted to 20cm
691B60052 | VISN 22: VA Desert Pacific Healthcare Network | 691 LOS ANGELES CA | XR PORT C ARM | Please adjust minimum vertical travel to 40 cm for fair and open competition | Adjusted to 40cm
691B60052 | VISN 22: VA Desert Pacific Healthcare Network | 691 LOS ANGELES CA | XR PORT C ARM | Please adjust minimum c-arm depth to 27" to allow for fair and open competition | Adjusted to 27"
691B60052 | VISN 22: VA Desert Pacific Healthcare Network | 691 LOS ANGELES CA | XR PORT C ARM | Please adjust minimum dual monitor requirement to 19" for fair and open competition | Adjusted to 19"
691B60052 | VISN 22: VA Desert Pacific Healthcare Network | 691 LOS ANGELES CA | XR PORT C ARM | Please adjust maximum matrix size to 2,048 x 2,048 for fair and open competition or remove | Adjusted to 2,048 x 2048
691B60052 | VISN 22: VA Desert Pacific Healthcare Network | 691 LOS ANGELES CA | XR PORT C ARM | Please adjust maximum dimensions for fair and open competition to: Length: 83" Height: 79" Width: 33" | Adjusted to length: 83, height: 79, width: 33
691B60052 | VISN 22: VA Desert Pacific Healthcare Network | 691 LOS ANGELES CA | XR PORT C ARM | Please adjust maximum weight for fair and open competition to 920 lbs | No, requested maximum weight remains 700lbs because increasing the weight significantly would introduce unnecessary mechanical bulk and reduce convenient mobility. You may still submit but this factor will affect consideration.
691B60052 | VISN 22: VA Desert Pacific Healthcare Network | 691 LOS ANGELES CA | XR PORT C ARM | Please remove UPS on board requirement for fair and open competition | Removed the onboard UPS requirement. Instead, battery backup power to ensure images are stored locally in the event of a power outage
691B60052 | VISN 22: VA Desert Pacific Healthcare Network | 691 LOS ANGELES CA | XR PORT C ARM | line 39 t. Can minimum C-arm mounted control monitor size be changed to 7.1 for fair competition? | Adjusted to 7.1
691B60052 | VISN 22: VA Desert Pacific Healthcare Network | 691 LOS ANGELES CA | XR PORT C ARM | line 103 h. can you please explain what grip refers to and/or measurements of grip? | With the flat detector lowered and the grip attached (do not count the grip on the measurement), the minimum distance from the detector to the source is 80 cm of free space.
691B60053 | VISN 22: VA Desert Pacific Healthcare Network | 691 LOS ANGELES CA | XR US | 2.b. Please remove navigation software; this is not applicable for this care area. | ED has interns and trainees who would utilize this mode to learn and thus provide better care for veterans. This requirement remains.
691B60053 | VISN 22: VA Desert Pacific Healthcare Network | 691 LOS ANGELES CA | XR US | 5.e. Please confirm if shear-wave and/or strain-based elastography is clinically needed for the ED. | Yes, shear-wave and strain-based elastography is clinically needed for the ED since it allows clinical users to view signs of heart attack earlier than an ECG would be able to detect.
693B63004 | VISN 4: VA Stars & Stripes Healthcare Network | 693 NEW ALLENTOWN PA | XR MAMMO | 1.d. Number of generators (3) - Please update to single-phase to align with industry standard. 1.h. SID 70cm - to allow for fair competition please update to 65-70 1.k. Maximum system weight [lbs]: 900 - Please increase to 925 lbs to allow for fair competition. What vaccum assisted device will you use for biopsy? Do you need a biopsy chair quoted? | 1.d. Ok to update to single Phase 1.h. Update to 45 - 70 cm 1.k. Update to 925lbs is ok 1.k. Not required - No niopsoes will be performed at this location where the system will be installed. No quote for chair is required
693B63005 | VISN 4: VA Stars & Stripes Healthcare Network | 693 WILKES BARRE PA | XR RAD | 2.k. Do you want full or partial system UPS? If full, please ensure site meets the dimensions and power specs. | One sized appropertatly to safely power down computers.
693B63005 | VISN 4: VA Stars & Stripes Healthcare Network | 693 WILKES BARRE PA | XR RAD | 3.q. Can you move Linear Tomography into the Value-Add section to increase competition? | no
693B63006 | VISN 4: VA Stars & Stripes Healthcare Network | 693 WILKES BARRE PA | XR US | 8.b. Can this Prostate Gland measurement be placed in the Value Add section? | No
693B63006 | VISN 4: VA Stars & Stripes Healthcare Network | 693 WILKES BARRE PA | XR US | 8.h. Can the Standard Planning Templates for prostate biopsies be added to the Value Add section? | no
693B63007 | VISN 4: VA Stars & Stripes Healthcare Network | 693 WILKES BARRE PA | XR RF | 1.C Is a third detector being asked for? If so, are there minimum specs? 7.D Is the “option for Automation” referring to the software automatically stitching images or is it referring to the equipment moving in an Automated manner to acquire multiple images to stitch? | 1.C No third detector is required , only 14x17 wireless detectors for the table and wall. QTY: 2 7.D Both software and Automatic Movement
693B63007 | VISN 4: VA Stars & Stripes Healthcare Network | 693 WILKES BARRE PA | XR RF | 1.ba. Can you changed the table detector to fixed detector? 3.u. Please remove bar code reader to allow for open competition. 6.a. Please change the control room monitor size to 23". 6.b. Please change the in-room monitor size to 23". 7.a. Please change move bone suppression to value adds? | 1.ba Detectors need to be wireless 3.u. Bar code reader is required to be aligned with the new EHRM requirements 6.a Monitor size can be changed to 23" 6.b Monitor size can be changed to 23" 7.a Bone Suppression is a requirement for clinical studies that will be performed.
756B65006 | VISN 17: VA Heart of Texas Health Care Network | 756 EL PASO TX | XR US | Column A does not include this Equipment PO - 756B65006 in order to pre populate the VISN, Station and Equipment type. Can you please verify the transducers that are being traded in with the system. Transducers need are not listed on matrix form. | Trade in information is already provided in the specification.
756B65007 | VISN 17: VA Heart of Texas Health Care Network | 756 EL PASO TX | XR PORT C-ARM | Would changing the minimum detector size to 30cm be considered a clinically equivalent or substantially compliant alternative to the specified 31cm minimum detector size? | The requirements will remain as is
756B65007 | VISN 17: VA Heart of Texas Health Care Network | 756 EL PASO TX | XR PORT C-ARM | Please adjust minimum requirement of panning to +/- 10 deg for fair and open competition | The requirements will remain as is
756B65007 | VISN 17: VA Heart of Texas Health Care Network | 756 EL PASO TX | XR PORT C-ARM | Please adjust minimum vertical travel to 40 cm for fair and open competition | The requirements will remain as is
756B65007 | VISN 17: VA Heart of Texas Health Care Network | 756 EL PASO TX | XR PORT C-ARM | Please adjust minimum c-arm depth to 27" to allow for fair and open competition | The requirements will remain as is
756B65007 | VISN 17: VA Heart of Texas Health Care Network | 756 EL PASO TX | XR PORT C-ARM | Please adjust minimum dual monitor requirement to 19" for fair and open competition | The requirements will remain as is
756B65007 | VISN 17: VA Heart of Texas Health Care Network | 756 EL PASO TX | XR PORT C-ARM | Please adjust the operator control monitor to 10" for fair and open competition | The requirements will remain as is
756B65007 | VISN 17: VA Heart of Texas Health Care Network | 756 EL PASO TX | XR PORT C-ARM | Please remove UPS on board requirement for fair and open competition | The requirements will remain as is
756B65007 | VISN 17: VA Heart of Texas Health Care Network | 756 EL PASO TX | XR PORT C-ARM | line 126. Serial number for second trade-in KS01934, can you please verify the serial number, this is listed as an OEC Elite CFD, this is not an OEC Elite serial number. | Serial number ES-1050 has been corrected on the technical specifications document
757B60015 | VISN 10: VA Healthcare System of Ohio | 757 COLUMBUS OH | XR US | Matrix 1.e. Can battery life be moved to Value Added section | Yes, Please see revised Tech Specs and Compliance Matrix for changes.